## DFS BUILD LIST
00224 | REV B | done
00128 | REV B | done
00129 | REV B | done
00216 | rev b | done
00313 | rev b | ASSIGN B

## Mounting Adhesive Calculation
bondline thickness calculation
[5 empty] | v | 0
radius | 12.7 | mm |  |  |  | 0.025
BK7 CTE | 7.1e-06 |  | Titanium CTE | 8.9e-06 |  | 0.05
Calcite Radial CTE | 8e-06 |  | SS CTE | 1.8e-05 |  | 0.075
Aluminum CTE | 2.1e-05 |  | RTV CTE | 3e-05 |  | 0.1
2216 CTE | 0.000102 |  |  |  |  | 0.125
| BK7 to aluminum 2216 |  |  |  |  | 0.15 | BK7 to aluminum RTV
| Bayar | Deluzio | Muench/Vanbezooijen |  |  | 0.175 | Bayar | Deluzio | Muench/Vanbezooijen
| 2.1794 | 2.1794 | 2.5013 |  |  | 0.2 | 19.6144 | 19.6144 | -123.8807
| 2.0472 | 2.0644 | 2.3427 |  |  | 0.225 | 16.751 | 17.9803 | -519.5977
| 1.9243 | 1.9558 | 2.1961 |  |  | 0.25 | 14.5198 | 16.5307 | 219.939
| 1.8098 | 1.8531 | 2.0603 |  |  | 0.275 | 12.7322 | 15.2359 | 87.9678
| 1.7029 | 1.7557 | 1.934 |  |  | 0.3 | 11.2679 | 14.0724 | 53.8566
| 1.6027 | 1.6634 | 1.8163 |  |  | 0.325 | 10.0464 | 13.0212 | 38.1981
| 1.5088 | 1.5756 | 1.7063 |  |  | 0.35 | 9.012 | 12.0668 | 29.2069
| 1.4205 | 1.4921 | 1.6034 |  |  | 0.375 | 8.1248 | 11.1964 | 23.3721
| 1.3373 | 1.4125 | 1.5068 |  |  | 0.4 | 7.3554 | 10.3994 | 19.2797
| 1.2589 | 1.3367 | 1.416 |  |  | 0.425 | 6.6818 | 9.6669 | 16.2507
| 1.1848 | 1.2642 | 1.3305 |  |  | 0.45 | 6.0872 | 8.9913 | 13.9183
| 1.1146 | 1.195 | 1.2498 |  |  | 0.475 | 5.5585 | 8.3664 | 12.0669
| 1.0481 | 1.1288 | 1.1735 |  |  | 0.5 | 5.0852 | 7.7865 | 10.5616
| 0.98498 | 1.0654 | 1.1013 |  |  |  | 4.6591 | 7.2469 | 9.3137
| 0.92499 | 1.0046 | 1.0329 |  |  |  | 4.2735 | 6.7437 | 8.2624
| 0.8679 | 0.94634 | 0.96798 |  |  |  | 3.9229 | 6.2733 | 7.3646
| 0.8135 | 0.89037 | 0.90625 |  |  |  | 3.6027 | 5.8325 | 6.589
| 0.76162 | 0.83658 | 0.84751 |  |  |  | 3.309 | 5.4186 | 5.9122
| 0.71208 | 0.78486 | 0.79153 |  |  |  | 3.0389 | 5.0293 | 5.3164
| 0.66472 | 0.73509 | 0.73814 |  |  |  | 2.7894 | 4.6625 | 4.788
| 0.6194 | 0.68715 | 0.68715 |  |  |  | 2.5584 | 4.3161 | 4.3161
Calcite to aluminum 2216
| Bayar | Deluzio | Muench/Vanbezooijen
| 2.0383 | 2.0383 | 2.3172
| 1.9146 | 1.9267 | 2.1727
| 1.7997 | 1.8217 | 2.0389
| 1.6926 | 1.7227 | 1.9147
| 1.5926 | 1.6293 | 1.7989
| 1.499 | 1.5409 | 1.6909
| 1.4111 | 1.4573 | 1.5897
| 1.3285 | 1.3779 | 1.4949
| 1.2508 | 1.3026 | 1.4059
| 1.1774 | 1.2309 | 1.322
| 1.1081 | 1.1627 | 1.2429
| 1.0424 | 1.0976 | 1.1682
| 0.98024 | 1.0356 | 1.0975
| 0.9212 | 0.97628 | 1.0306
| 0.86509 | 0.91958 | 0.96702
| 0.8117 | 0.8653 | 0.90664
| 0.76083 | 0.8133 | 0.84921
| 0.71231 | 0.76343 | 0.7945
| 0.66597 | 0.71556 | 0.74233
| 0.62168 | 0.66958 | 0.69252
| 0.5793 | 0.62538 | 0.64492
BK7 to SS 2216
| Bayar | Deluzio | Muench/Vanbezooijen
| 1.648 | 1.648 | 1.8257
| 1.5514 | 1.5626 | 1.717
| 1.4612 | 1.4819 | 1.6158
| 1.3769 | 1.4053 | 1.5213
| 1.2978 | 1.3326 | 1.4329
| 1.2235 | 1.2635 | 1.3499
| 1.1536 | 1.1978 | 1.2719
| 1.0877 | 1.1352 | 1.1984
| 1.0254 | 1.0754 | 1.1292
| 0.96652 | 1.0183 | 1.0637
| 0.91072 | 0.96377 | 1.0018
| 0.85779 | 0.91157 | 0.94308
| 0.80751 | 0.86157 | 0.88737
| 0.75968 | 0.81364 | 0.83443
| 0.71412 | 0.76764 | 0.78405
| 0.67069 | 0.72348 | 0.73606
| 0.62923 | 0.68103 | 0.69028
| 0.58961 | 0.6402 | 0.64657
| 0.55171 | 0.6009 | 0.6048
| 0.51543 | 0.56304 | 0.56483
| 0.48066 | 0.52655 | 0.52655
Calcite to SS 2216
| Bayar | Deluzio | Muench/Vanbezooijen
| 1.5119 | 1.5119 | 1.6601
| 1.4233 | 1.434 | 1.563
| 1.3406 | 1.3602 | 1.4722
| 1.2632 | 1.2902 | 1.3874
| 1.1906 | 1.2238 | 1.3078
| 1.1225 | 1.1606 | 1.233
| 1.0583 | 1.1004 | 1.1626
| 0.99786 | 1.0431 | 1.0963
| 0.94074 | 0.98833 | 1.0336
| 0.88671 | 0.93604 | 0.97426
| 0.83553 | 0.88605 | 0.91807
| 0.78697 | 0.83819 | 0.86476
| 0.74083 | 0.79234 | 0.8141
| 0.69695 | 0.74836 | 0.76591
| 0.65516 | 0.70616 | 0.72002
| 0.61531 | 0.66562 | 0.67625
| 0.57727 | 0.62664 | 0.63447
| 0.54093 | 0.58915 | 0.59455
| 0.50616 | 0.55305 | 0.55635
| 0.47287 | 0.51827 | 0.51978
| 0.44097 | 0.48473 | 0.48473
BK7 to Titanium 2216
nu | Bayar | Deluzio | Muench/Vanbezooijen
0 | 0.28222 | 0.24554 | 0.32391 | 0.012752
0.025 | 0.2651 | 0.23345 | 0.30337 | 0.011944
0.05 | 0.24919 | 0.22195 | 0.28439 | 0.011197
0.075 | 0.23436 | 0.211 | 0.2668 | 0.010504
0.1 | 0.22051 | 0.20055 | 0.25044 | 0.00986
0.125 | 0.20755 | 0.19057 | 0.2352 | 0.0092599
0.15 | 0.19538 | 0.18103 | 0.22096 | 0.0086994
0.175 | 0.18395 | 0.17191 | 0.20763 | 0.0081745
0.2 | 0.17318 | 0.16317 | 0.19512 | 0.007682
0.225 | 0.16302 | 0.15479 | 0.18337 | 0.0072191
0.25 | 0.15342 | 0.14676 | 0.17229 | 0.0067831
0.275 | 0.14434 | 0.13904 | 0.16184 | 0.0063717
0.3 | 0.13573 | 0.13163 | 0.15197 | 0.0059829
0.325 | 0.12755 | 0.1245 | 0.14262 | 0.0056149
0.35 | 0.11978 | 0.11763 | 0.13376 | 0.0052661
0.375 | 0.11239 | 0.11102 | 0.12535 | 0.004935
0.4 | 0.10535 | 0.10465 | 0.11736 | 0.0046203
0.425 | 0.098627 | 0.09851 | 0.10975 | 0.0043208
0.45 | 0.092211 | 0.092581 | 0.1025 | 0.0040355
0.475 | 0.086079 | 0.086856 | 0.095587 | 0.0037633
0.5 | 0.080211 | 0.081323 | 0.088984 | 0.0035033

## Ideal 1 calcite piece
Calcite properties
88190000000 | N/m^2 | E in X | 80000 | N/m^2 | stress upper limit
72000000000 | N/m^2 | E in Y | 35 | C | temp low | 308.15 | K
72000000000 | N/m^2 | E in Z | 45 | C | temp high | 318.15 | K
2.81 | kg/m^3 | density | 10 | degrees | delta T
2.5e-05 | per Kelvin | Thermal exp X | 28.162 | mm | diameter
5.8e-06 | per Kelvin | Thermal exp Y | 2.7 | mm | thickness
5.8e-06 | per Kelvin | Thermal exp Z
5.526 | W/(m*K) | thermal cond X
4.646 | W/(m*K) | thermal cond Y
4.646 | W/(m*K) | thermal cond Z
834.3 | J/(kg*K) | specific heat
growth from delta T |  |  | FEA
0.0070405 | mm | FEA looks good | 0.00715 | mm
0.0016334 | mm | FEA looks good | 0.00161 | mm
0.0001566 | mm | FEA is a bit off | 0.000306 | mm
Stress in a floating example is well below, and now we have an idea of the starting point.

## Energy Hitting FiltersWindow
thuill_astme590_lambda.png (1280×640) (nasa.gov)
100mm diameter objective lens
7853.9816 | mm^2 | Area on face of lens, adding 10% for safety
1366 W/m2 | NASA - SOLAR IRRADIANCE
0.007854 | m^2 | Area converter to multiply by irradiance
10.7285 | Watts
25 | mm | diameter of beam hitting filters
490.8739 | mm^2 | area of beam
0.00049087 | m^2 | area of beam
Heat rejection filter
650nm CWL, 50mm Dia. Hard Coated OD 4.0 50nm Bandpass Filter (edmundoptics.com)
625 | 675 | 50 | nm width | 1 micron = 1000 nanometer
using 1000 W/m^2/um to be safe
effectively, the filter allows 1 W/m^2/nm
1
0.95 | transmission of th 50 nm width
[5 empty] | kW = (WT x Cp x Δ T)/3412 x h |  |  | Δ T | 3447.1102
heat rejection lets how many Watts through? |  |  |  |  | (3412 x h *kW)/(WT x Cp) = Δ T)
0.37306 |  |  |  |  | Where:
kW = your kilowatt requirement
heat rejection absorbs? |  |  |  |  | WT= the weight of the material to be heated, in lbs.
10.3555 | Watts |  | 21096 |  | Cp = the specific heat of the material to be heated, in BTU/lb°F
Δ T = Temperature Rise, in °F
3412 = Conversion Factor, BTU/kWh
range of prefilter <1 nm |  |  |  |  | h = how long it should take to reach the temperature set point, in hours
approximately calculating Wattage
0.0037306 | Watts allowed through |  |  |  | WT | 0.05
[5 empty] | Cp | 0.205
prefilter absorbs |  |  |  |  | factor | 3412
0.36933 |  |  |  |  | h | 1
Lyot window Wattage estimate we will use 0.1 W just to be safe

## Lyot Internal Heat Flow
Estimates of Radiation vs Conduction vs Convection from heaters to the optics in the Lyot filter
Radiation driving equation
| T1 | 35 | C
| T2 | 0 | C | starting from coldest
| ϵ1 | 0.9 |  |  | Emissitivity of Common Materials (omega.co.uk)
| ϵ2 | 0.9
| σ | 5.67e-08 | W/(m^2*K^4)
| T1 | 308.15 | K
| T2 | 273.15 | K
exposed optic to heater area |  | 0.000228 | m^2 | 4 of the drawing's areas
exposed to all hot elements |  | 0.03 | m^2
| q rate | 0.03649 | Watts | just optic to heater
| q rate | 4.8014 | Radiation Watts | all surfaces
Conduction driving equations
| TC alum | 167 | W/(m*K)
| TC 2216 | 0.395 | W/(m*K)
| TC BK7 | 1.114 | W/(m*K)
| TC series | 168.509 | W/(m*K)
| T1 | 308.15 | K
| T2 | 273.15 | K
| total area | 0.0007539 | m^2
| t of 2216 | 0.002 | m
| t of BK7 | 0.0508 | m
| t of Al | 0.01 | m
| total t | 0.0628 | m
| q rate | 70.802 | Conduction Watts
Convection driving equations
| Air h | 3 | W/(m^2 * K)
exposed to all hot elements |  | 0.03 | m^2
| T1 | 308.15 | K
| T2 | 273.15 | K
| q | 3.15 | Convection Watts
Heater Total Wattage Estimate
| Heater | 78.7533 | Watts
𝑞 ̇=𝜎 ((𝑇_1^4−𝑇_2^4 ))/(1⁄𝜖_1 + 1⁄𝜖_2 −1)

## spring plunger calc
Lyot Filter Design
mix of classwork and KSO final, calculations are good | [10 empty] | Element | Material | Diameter | Thickness
delta T for stack survival | 55 | C
name | material | thickness | stiffness (N/mm) | 1/stiffness | CTE axial | Growth by delta T (mm) | [6 empty] | (mm) | (mm)
from FEA
gasket |  |  | 430 | 0.0023256
window | N-BK7 | 10 | 16345210.8532 | 6.118e-08 | 8.3e-06 | 0.004565
polarizer | N-BK7 | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
quarter-waveplate |  | 13 | 11528706.4791 | 8.674e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 2.7 | 23169601.4829 | 4.316e-08 | 3e-05 | 0.004455
half-waveplate |  | 13 | 11528706.4791 | 8.674e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 2.7 | 23169601.4829 | 4.316e-08 | 3e-05 | 0.004455
polarizer | N-BK7 | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 10.8 | 6049606.7756 | 1.653e-07 | 3e-05 | 0.01782
half-waveplate |  | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 10.8 | 6049606.7756 | 1.653e-07 | 3e-05 | 0.01782
polarizer | N-BK7 | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 5.4 | 14806040.8647 | 6.754e-08 | 3e-05 | 0.00891
half-waveplate |  | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 5.4 | 14806040.8647 | 6.754e-08 | 3e-05 | 0.00891
quarter-waveplate |  | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
half-waveplate |  | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
polarizer | N-BK7 | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 1.35 | 59101654.8463 | 1.692e-08 | 3e-05 | 0.0022275
half-waveplate |  | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
crystal | calcite | 1.35 | 59101654.8463 | 1.692e-08 | 3e-05 | 0.0022275
quarter-waveplate |  | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
polarizer | N-BK7 | 13 | 11534025.3749 | 8.67e-08 | 8.3e-06 | 0.0059345
window |  | 10 | 16345210.8532 | 6.118e-08 | 8.3e-06 | 0.004565
gasket |  |  | 430 | 0.0023256
Optics cap | delrin | 5 | 20000 | 5e-05 |  |  |  | come back to calculate thickness of whole stack
0.0027837
Totals |  |  | 388896351.727 | 212.6303 |  | 0.1531
total axial stiffness | 212.6303 | N/mm | stiffness of 1 aluminum rib | 66 | N/mm | stiffness of 1 steel rib | 183.5536 | N/mm
|  |  | 4 aluminum ribs together | 264 | N/mm | 4 steel ribs together | 734.2144 | N/mm | this makes sense, CTE Al is 21, CTE Steel of 16
|  |  | axial growth of rib if aluminum | 0.28 | mm | axial growth of rib if steel | 0.2 | mm | fea has this axial growths slightly worse than 1 D calculation
growth from z estimate 55 C delta T | 0.1531 | mm |  |  |  |  | working on matching this exactly
axial growth of rib if steel | 0.2 | mm |  |  |  | this is done
difference | 0.046896 | mm
preload from 55 C delta T | 9.9716 | N | 2.2436 |  | lbs
margin for index m fluid | 2 | N | 0.45 |  | lbs
preload for assembly analysis | 11.9716 | N | 2.6936 |  | lbs
per spring 4 total | 1.5 mm compression engagement
| 1.5 | mm
| 0.059055 | in
parrallel stiffnes of optic stack and ribs | 946.8447 | N/mm
preload taking ribs into account | 44.4037 | N
total mass of optics and mounts | 1600 | grams
force due to gravity of optics and mounts | 16 | N
total preload needed | 60.4037 | N
minimum force per spring plunger | 15.1009 | N
total preload needed | 13.5908 | lbs | this case is will a full delta T of 55 and all of the optics weight is facing toward the window
minimum force per spring plunger | 3.3977 | lbs
k needed | 40.2691 | N/mm
k needed | 230.1381 | N/in
per spring k needed | 10.0673 | N/mm | 2.2651 | lb/mm
per spring k needed | 57.5345 | N/in | 12.9453 | lb/in
8969T604 k value | 2.25 | lb/mm | 57.15
little high but very close
max load 4 springs | 19.44 | lbs
length of spring | 9.7 | mm
compression | 1.5 | mm
extra | 0.5 | mm
into optics cap | 3 | mm
into entrance cap | 4.7 | mm
[6 empty] | 27838541 | (N/m)/m^2
Note on A45: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: include a preload to compress the o ring 7%

## Optic Cap FEA
Purpose is to get rough stiffness number to add to the next tab
holds the springs, rubber gasket, and insulates mostly
Force | 1 | N
disp | 5e-05 | mm
stiffness | 20000 | N/mm

## Calcite jig tolerance
diameter | 60 | mm
typ allow | 0.05 | mm
opp/adj | 0.0016667
tan^-1 | 0.095493 | deg
per calcite this is the allowed tilt

## Index Matching Gel
|  |  | index of material | index of OC 440
3 | window
4 | polarizer | BK7 | 1.5
5 | quarter-waveplate | BK7 | 1.5
6 | crystal | calcite | 1.5
7 | half-waveplate | BK7 | 1.5
8 | crystal | calcite | 1.5
9 | polarizer | BK7 | 1.5
10 | crystal | calcite | 1.5
11 | half-waveplate | BK7 | 1.5
12 | crystal | calcite | 1.5
13 | polarizer | BK7 | 1.5
14 | crystal | calcite | 1.5
15 | half-waveplate | BK7 | 1.5
16 | crystal | calcite | 1.5
17 | quarter-waveplate | BK7 | 1.5
18 | half-waveplate | BK7 | 1.5
19 | polarizer | BK7 | 1.5
20 | crystal | calcite | 1.5
21 | half-waveplate | BK7 | 1.5
22 | crystal | calcite | 1.5
23 | quarter-waveplate | BK7 | 1.5
24 | polarizer | BK7 | 1.5
25 | window

## Offset is better
Calcite Part values only, Offset adhesive is better overall
| Aligned Adhesive | Offset by 60 Deg |  | Difference from Algined to Offset
X Normal Mpa | 0.02123 | 0.02326 |  | -0.00203
Y Normal Mpa | 0.01635 | 0.01585 |  | 0.0005
XY Shear Mpa | 0.06561 | 0.06542 |  | 0.00019
YZ Shear Mpa | 0.05671 | 0.05575 |  | 0.00096
X Disp mm | 0.08584 | 0.08584 |  | 0
Y Disp mm | -0.004617 | 0.04648 |  | -0.051097
Optimizing the Offset
Design Study 1
| Scenarios/Iterations: | 98
| Parameter Constraint or Goal | Format | Unit | Initial Value | Optimal Value | Scenario 1 | Scenario 2 | Scenario 3 | Scenario 4 | Scenario 5 | Scenario 6 | Scenario 7 | Scenario 8 | Scenario 9 | Scenario 10 | Scenario 11 | Scenario 12 | Scenario 13 | Scenario 14 | Scenario 15 | Scenario 16 | Scenario 17 | Scenario 18 | Scenario 19 | Scenario 20 | Scenario 21 | Scenario 22 | Scenario 23 | Scenario 24 | Scenario 25 | Scenario 26 | Scenario 27 | Scenario 28 | Scenario 29 | Scenario 30 | Scenario 31 | Scenario 32 | Scenario 33 | Scenario 34 | Scenario 35 | Scenario 36 | Scenario 37 | Scenario 38 | Scenario 39 | Scenario 40 | Scenario 41 | Scenario 42 | Scenario 43 | Scenario 44 | Scenario 45 | Scenario 46 | Scenario 47 | Scenario 48 | Scenario 49 | Scenario 50 | Scenario 51 | Scenario 52 | Scenario 53 | Scenario 54 | Scenario 55 | Scenario 56 | Scenario 57 | Scenario 58 | Scenario 59 | Scenario 60 | Scenario 61 | Scenario 62 | Scenario 63 | Scenario 64 | Scenario 65 | Scenario 66 | Scenario 67 | Scenario 68 | Scenario 69 | Scenario 70 | Scenario 71 | Scenario 72 | Scenario 73 | Scenario 74 | Scenario 75 | Scenario 76 | Scenario 77 | Scenario 78 | Scenario 79 | Scenario 80 | Scenario 81 | Scenario 82 | Scenario 83 | Scenario 84 | Scenario 85 | Scenario 86 | Scenario 87 | Scenario 88 | Scenario 89 | Scenario 90 | Scenario 91 | Scenario 92 | Scenario 93 | Scenario 94 | Scenario 95 | Scenario 96 | Scenario 97 | Scenario 98
[6 empty] | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated
| Ring Outer Dia |  | mm | 38.5 | 38.5 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45
| Ring Inner Dia |  | mm | 35 | 37.5 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 35.5 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 36.5 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 37.5 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38 | 38
| YZ Shear Cal | Minimize | N/mm^2 | 0.048718 | 0.013602 | 0.048718 | 0.050418 | 0.05256 | 0.053077 | 0.054322 | 0.054609 | 0.054646 | 0.054949 | 0.055718 | 0.055579 | 0.055492 | 0.055617 | 0.056187 | 0.055747 | 0.034824 | 0.036112 | 0.036791 | 0.037565 | 0.038452 | 0.038958 | 0.03931 | 0.039342 | 0.039523 | 0.039621 | 0.039324 | 0.039153 | 0.03964 | 0.039247 | 0.022235 | 0.024207 | 0.024052 | 0.0245 | 0.025717 | 0.025618 | 0.025357 | 0.025793 | 0.026639 | 0.026529 | 0.026741 | 0.026124 | 0.027034 | 0.026868 | 0.016727 | 0.01799 | 0.018698 | 0.018663 | 0.019343 | 0.019504 | 0.020029 | 0.020023 | 0.02023 | 0.020456 | 0.020429 | 0.020124 | 0.02074 | 0.020469 | 0.013078 | 0.01394 | 0.014635 | 0.015033 | 0.015407 | 0.01557 | 0.015752 | 0.015887 | 0.015936 | 0.016096 | 0.016015 | 0.016225 | 0.01641 | 0.016863 | 0.013602 | 0.017111 | 0.019375 | 0.020776 | 0.021801 | 0.0218 | 0.022307 | 0.022471 | 0.022941 | 0.022785 | 0.023114 | 0.023061 | 0.023211 | 0.023423 | 0.012006 | 0.012412 | 0.014805 | 0.016467 | 0.017774 | 0.018283 | 0.018684 | 0.018889 | 0.019181 | 0.019323 | 0.019407 | 0.01938 | 0.019774 | 0.019706
| XY Shear Cal | Minimize | N/mm^2 | 0.055265 | 0.022827 | 0.055265 | 0.058002 | 0.05943 | 0.061067 | 0.061631 | 0.061924 | 0.06371 | 0.063731 | 0.063381 | 0.063845 | 0.064198 | 0.064872 | 0.064889 | 0.065425 | 0.039939 | 0.0415 | 0.042681 | 0.043952 | 0.044988 | 0.04478 | 0.045353 | 0.045224 | 0.04639 | 0.046543 | 0.046773 | 0.046442 | 0.047875 | 0.047278 | 0.036215 | 0.038467 | 0.039397 | 0.04038 | 0.04257 | 0.042388 | 0.04236 | 0.042933 | 0.044532 | 0.044494 | 0.044452 | 0.044051 | 0.04617 | 0.046457 | 0.031845 | 0.035236 | 0.03625 | 0.036883 | 0.039626 | 0.039721 | 0.040276 | 0.040152 | 0.040969 | 0.041333 | 0.041322 | 0.040611 | 0.042904 | 0.042985 | 0.027994 | 0.033124 | 0.035149 | 0.036096 | 0.038233 | 0.038981 | 0.039777 | 0.040064 | 0.041231 | 0.040888 | 0.040457 | 0.040886 | 0.042004 | 0.042215 | 0.022827 | 0.029445 | 0.03317 | 0.035654 | 0.037692 | 0.038636 | 0.039087 | 0.039355 | 0.041032 | 0.040506 | 0.041167 | 0.041212 | 0.041907 | 0.042803 | 0.019715 | 0.024654 | 0.029962 | 0.033551 | 0.036867 | 0.037889 | 0.039206 | 0.039138 | 0.040369 | 0.040627 | 0.040872 | 0.041296 | 0.042475 | 0.042278
| Y Normal Cal Stress | Minimize | N/mm^2 | 0.012413 | 0.023043 | 0.012413 | 0.012706 | 0.013346 | 0.013755 | 0.014006 | 0.01417 | 0.014535 | 0.014565 | 0.01457 | 0.015029 | 0.014982 | 0.01557 | 0.015479 | 0.015855 | 0.012467 | 0.012619 | 0.012741 | 0.01289 | 0.01343 | 0.013453 | 0.013467 | 0.013664 | 0.01446 | 0.014744 | 0.014313 | 0.015028 | 0.015055 | 0.016013 | 0.013863 | 0.014251 | 0.013072 | 0.012413 | 0.014697 | 0.014204 | 0.014178 | 0.014036 | 0.015837 | 0.015807 | 0.015537 | 0.015031 | 0.017133 | 0.017057 | 0.017496 | 0.017773 | 0.01693 | 0.016003 | 0.017599 | 0.017954 | 0.01819 | 0.017753 | 0.018594 | 0.018706 | 0.018787 | 0.01748 | 0.02009 | 0.019825 | 0.017937 | 0.017219 | 0.016826 | 0.015887 | 0.017057 | 0.016845 | 0.017252 | 0.017344 | 0.018211 | 0.018724 | 0.017224 | 0.017886 | 0.018992 | 0.018861 | 0.023043 | 0.019228 | 0.017043 | 0.017121 | 0.018094 | 0.016834 | 0.017502 | 0.017235 | 0.019183 | 0.018565 | 0.018437 | 0.018751 | 0.019275 | 0.019671 | 0.024078 | 0.023465 | 0.018901 | 0.016982 | 0.017848 | 0.017254 | 0.017893 | 0.017078 | 0.018553 | 0.018417 | 0.018195 | 0.01812 | 0.019568 | 0.019627
| X Normal Stress | Minimize | N/mm^2 | 0.017952 | 0.026304 | 0.017952 | 0.01757 | 0.01923 | 0.020485 | 0.018563 | 0.019157 | 0.019526 | 0.021007 | 0.019478 | 0.020012 | 0.021399 | 0.023358 | 0.021142 | 0.02326 | 0.015383 | 0.015205 | 0.01655 | 0.017471 | 0.016186 | 0.016813 | 0.017184 | 0.018718 | 0.017581 | 0.017894 | 0.019311 | 0.020151 | 0.01897 | 0.022112 | 0.022165 | 0.020657 | 0.021746 | 0.022313 | 0.020707 | 0.021759 | 0.020718 | 0.021264 | 0.021504 | 0.023173 | 0.024183 | 0.023436 | 0.023073 | 0.02469 | 0.025807 | 0.024154 | 0.025793 | 0.026351 | 0.023911 | 0.024472 | 0.026304 | 0.027418 | 0.024312 | 0.026738 | 0.027587 | 0.030291 | 0.027613 | 0.030406 | 0.028057 | 0.027264 | 0.026854 | 0.028247 | 0.027284 | 0.027098 | 0.028097 | 0.029038 | 0.026714 | 0.029257 | 0.029903 | 0.031879 | 0.030499 | 0.032367 | 0.026304 | 0.024255 | 0.024328 | 0.025049 | 0.02287 | 0.023092 | 0.024325 | 0.025311 | 0.022815 | 0.025062 | 0.025971 | 0.026354 | 0.025794 | 0.028124 | 0.033323 | 0.025328 | 0.024443 | 0.023754 | 0.022281 | 0.023387 | 0.023508 | 0.024174 | 0.02339 | 0.025128 | 0.024639 | 0.026717 | 0.024868 | 0.026888

## Bondline Thicknesses vs Theory
Optimization Numbers (3D effects, 2 Aluminum Parts vs 2D and 1 solid metal part)
Optic | Material | Bondline Thickness (mm) | v = 0.43 | % from 2D theory
Polarizer/Retarders | N-BK7 | 0.6 | 1.42 | 42.25%
Calcite 1.35 | Calcite | 1 | 0.66 | 51.52%
Calcite 2.7 | Calcite | 1 | 0.66 | 51.52%
Calcite 5.4 | Calcite | 1 | 0.66 | 51.52%
Calcite 10.8 | Calcite | 1 | 0.66 | 51.52%
2.84
53.64
n

## 2 thin Poker Optimization
| 2216 | 27838541 | (N/m) / (m^2)
| 2216 ring | 2783854.1 | (N/m) / (m^2)
| Elastic Support | 0.43 | N/m
doyle-2002.pdf (arizona.edu)
doyle-2002.pdf (arizona.edu)
https://wp.optics.arizona.edu/optomech/wp-content/uploads/sites/53/2016/11/423-Final-Report-Armstrong.pdf

## Linear Polarizer Analysis
adhesive thickness to optic | 2.5 | mm
inner ring inner dia | 55.5 | mm
inner ring outer dia | 67 | mm
Design Project Requirements Review (arizona.edu)
2216 poisson ratio | 0.43
from AAC report, essentially reran with the KSO optic and overlaping chips, and 0-30 instead of 0-50

## KSO Thermal Assembly Results
KSO 1 80W heater 30 C goal
Goal Name | Unit | Value | Averaged Value | Minimum Value | Maximum Value | Progress [%] | Use In Convergence | Delta | Criteria
Front Optic Min Temp | [°C] | 18.28 | 18.28 | 18.28 | 18.29 | 64 | Yes | 0.0097 | 0.00621 | [5 empty] | 1 80 W Heater | 3 26 W Heaters | 3 15 W Heaters
Front Optic Avg Temp | [°C] | 25.98 | 25.99 | 25.98 | 26 | 47 | Yes | 0.01 | 0.00643 |  |  |  |  | Min Min Temp | 18.28 | 29.88 | 20.35
Front Optic Max Temp | [°C] | 29.67 | 29.68 | 29.67 | 29.69 | 83 | Yes | 0.02 | 0.01 |  |  |  |  | Max Max Temp | 30.87 | 43.46 | 31.01
Mid Optic Min Temp | [°C] | 29.61 | 29.61 | 29.61 | 29.61 | 89 | Yes | 0.00773 | 0.00687
Mid Optic Avg Temp | [°C] | 30.43 | 30.44 | 30.43 | 30.45 | 46 | Yes | 0.02 | 0.00718
Mid Optic Max Temp | [°C] | 30.85 | 30.86 | 30.85 | 30.87 | 64 | Yes | 0.02 | 0.01
Rear Optic Min Temp | [°C] | 18.59 | 18.58 | 18.57 | 18.59 | 46 | Yes | 0.02 | 0.01
Rear Optic Avg Temp | [°C] | 27.38 | 27.39 | 27.38 | 27.4 | 46 | Yes | 0.02 | 0.0074
Rear Optic Max Temp | [°C] | 30.53 | 30.54 | 30.53 | 30.55 | 100 | Yes | 0.02 | 0.05
KSO 3 26 W heaters 30 C goal
Goal Name | Unit | Value | Averaged Value | Minimum Value | Maximum Value | Progress [%] | Use In Convergence | Delta | Criteria
Front Optic Min Temp | [°C] | 29.88 | 29.88 | 29.88 | 29.88 | 100 | Yes | 6.05e-05 | 0.00284
Front Optic Avg Temp | [°C] | 39.21 | 39.21 | 39.21 | 39.21 | 100 | Yes | 7.46e-05 | 0.00372
Front Optic Max Temp | [°C] | 43.28 | 43.28 | 43.28 | 43.28 | 100 | Yes | 0.000149 | 0.01
Mid Optic Min Temp | [°C] | 42.01 | 42.01 | 42.01 | 42.01 | 100 | Yes | 5.31e-05 | 0.00402
Mid Optic Avg Temp | [°C] | 42.82 | 42.82 | 42.82 | 42.82 | 100 | Yes | 0.000279 | 0.00409
Mid Optic Max Temp | [°C] | 43.46 | 43.46 | 43.46 | 43.46 | 100 | Yes | 0.000438 | 0.01
Rear Optic Min Temp | [°C] | 29.91 | 29.91 | 29.91 | 29.92 | 44 | Yes | 0.01 | 0.00461
Rear Optic Avg Temp | [°C] | 39.09 | 39.09 | 39.09 | 39.09 | 100 | Yes | 0.00119 | 0.00384
Rear Optic Max Temp | [°C] | 42.23 | 42.23 | 42.21 | 42.24 | 100 | Yes | 0.02 | 0.05
KSO 3 15 W heaters 30 C goal
Goal Name | Unit | Value | Averaged Value | Minimum Value | Maximum Value | Progress [%] | Use In Convergence | Delta | Criteria
Front Optic Min Temp | [°C] | 20.35 | 20.35 | 20.35 | 20.35 | 100 | Yes | 0.00204 | 0.00265
Front Optic Avg Temp | [°C] | 27.08 | 27.08 | 27.07 | 27.08 | 100 | Yes | 0.00344 | 0.00355
Front Optic Max Temp | [°C] | 30.37 | 30.36 | 30.36 | 30.37 | 100 | Yes | 0.0069 | 0.01
Mid Optic Min Temp | [°C] | 30.09 | 30.09 | 30.08 | 30.09 | 82 | Yes | 0.00486 | 0.00401
Mid Optic Avg Temp | [°C] | 30.62 | 30.62 | 30.62 | 30.62 | 57 | Yes | 0.00726 | 0.00416
Mid Optic Max Temp | [°C] | 31.01 | 31.01 | 31 | 31.01 | 95 | Yes | 0.00986 | 0.00937
Rear Optic Min Temp | [°C] | 21.96 | 21.96 | 21.95 | 21.96 | 54 | Yes | 0.0057 | 0.00307
Rear Optic Avg Temp | [°C] | 28.29 | 28.29 | 28.29 | 28.29 | 75 | Yes | 0.00513 | 0.00384
Rear Optic Max Temp | [°C] | 30.7 | 30.7 | 30.69 | 30.7 | 100 | Yes | 0.00871 | 0.04
3 26 W heaters -10 C to 45 C
|  |  |  | Goal Name | Unit | Value | Averaged Value | Minimum Value | Maximum Value | Progress [%] | Use In Convergence | Delta | Criteria
|  |  |  | Front Optic Min Temp | [°C] | 27.95 | 27.95 | 27.95 | 27.95 | 100 | Yes | 0.00164 | 0.00375
|  |  |  | Front Optic Avg Temp | [°C] | 39.7 | 39.7 | 39.7 | 39.7 | 100 | Yes | 0.00211 | 0.00494
|  |  |  | Front Optic Max Temp | [°C] | 45.12 | 45.12 | 45.12 | 45.12 | 100 | Yes | 0.000772 | 0.02
|  |  |  | Mid Optic Min Temp | [°C] | 44.83 | 44.84 | 44.83 | 44.84 | 100 | Yes | 0.0022 | 0.00558
|  |  |  | Mid Optic Avg Temp | [°C] | 45.38 | 45.38 | 45.38 | 45.38 | 100 | Yes | 0.000476 | 0.00577
|  |  |  | Mid Optic Max Temp | [°C] | 45.73 | 45.73 | 45.73 | 45.73 | 100 | Yes | 0.00424 | 0.01
|  |  |  | Rear Optic Min Temp | [°C] | 30.38 | 30.38 | 30.38 | 30.38 | 100 | Yes | 0.000417 | 0.00435
|  |  |  | Rear Optic Avg Temp | [°C] | 41.64 | 41.64 | 41.64 | 41.64 | 100 | Yes | 0.00112 | 0.00533
|  |  |  | Rear Optic Max Temp | [°C] | 45.27 | 45.27 | 45.27 | 45.27 | 100 | Yes | 0.00325 | 0.06
3 40 W heater -10 C to 45 C
|  |  |  | Goal Name | Unit | Value | Averaged Value | Minimum Value | Maximum Value | Progress [%] | Use In Convergence | Delta | Criteria
|  |  |  | Front Optic Min Temp | [°C] | 40.5 | 40.5 | 40.5 | 40.5 | 100 | Yes | 1.24e-05 | 0.00377
|  |  |  | Front Optic Avg Temp | [°C] | 56.04 | 56.04 | 56.04 | 56.04 | 100 | Yes | 1.73e-05 | 0.00493
|  |  |  | Front Optic Max Temp | [°C] | 62.88 | 62.88 | 62.88 | 62.88 | 100 | Yes | 1.59e-05 | 0.02
|  |  |  | Mid Optic Min Temp | [°C] | 62.33 | 62.33 | 62.33 | 62.33 | 100 | Yes | 1.65e-05 | 0.00541
|  |  |  | Mid Optic Avg Temp | [°C] | 62.78 | 62.78 | 62.78 | 62.78 | 100 | Yes | 1.67e-05 | 0.00544
|  |  |  | Mid Optic Max Temp | [°C] | 63.09 | 63.09 | 63.09 | 63.09 | 100 | Yes | 3.31e-05 | 0.01
|  |  |  | Rear Optic Min Temp | [°C] | 43.83 | 43.83 | 43.83 | 43.83 | 100 | Yes | 8.84e-06 | 0.00405
|  |  |  | Rear Optic Avg Temp | [°C] | 58.12 | 58.12 | 58.12 | 58.12 | 100 | Yes | 1.36e-05 | 0.00509
|  |  |  | Rear Optic Max Temp | [°C] | 62.72 | 62.72 | 62.72 | 62.72 | 100 | Yes | 1.42e-05 | 0.07
|  |  |  | Goal Name | Unit | Value | Averaged Value | Minimum Value | Maximum Value | Progress [%] | Use In Convergence | Delta | Criteria
|  |  |  | Front Optic Min Temp | [°C] | 38.8 | 38.8 | 38.8 | 38.8 | 100 | Yes | 0.000232 | 0.00371
|  |  |  | Front Optic Avg Temp | [°C] | 45.49 | 45.49 | 45.49 | 45.49 | 100 | Yes | 0.000377 | 0.00433
|  |  |  | Front Optic Max Temp | [°C] | 46.73 | 46.73 | 46.73 | 46.73 | 100 | Yes | 0.000404 | 0.01
|  |  |  | Mid Optic Min Temp | [°C] | 45.53 | 45.53 | 45.53 | 45.53 | 100 | Yes | 0.000408 | 0.00454
|  |  |  | Mid Optic Avg Temp | [°C] | 46.21 | 46.21 | 46.21 | 46.21 | 100 | Yes | 0.000399 | 0.00456
|  |  |  | Mid Optic Max Temp | [°C] | 46.7 | 46.7 | 46.7 | 46.7 | 100 | Yes | 0.000404 | 0.00468
|  |  |  | Rear Optic Min Temp | [°C] | 31.01 | 31.01 | 31.01 | 31.01 | 100 | Yes | 4.03e-05 | 0.00332
|  |  |  | Rear Optic Avg Temp | [°C] | 42.13 | 42.13 | 42.13 | 42.13 | 100 | Yes | 0.000381 | 0.00425
|  |  |  | Rear Optic Max Temp | [°C] | 45.9 | 45.9 | 45.9 | 45.9 | 100 | Yes | 0.000397 | 0.01

## Calcite 5_4 optimization
Design Study 1
Scenarios/Iterations: | 378
Parameter Constraint or Goal | Format | Unit | Initial Value | Optimal Value | Scenario 1 | Scenario 2 | Scenario 3 | Scenario 4 | Scenario 5 | Scenario 6 | Scenario 7 | Scenario 8 | Scenario 9 | Scenario 10 | Scenario 11 | Scenario 12 | Scenario 13 | Scenario 14 | Scenario 15 | Scenario 16 | Scenario 17 | Scenario 18 | Scenario 19 | Scenario 20 | Scenario 21 | Scenario 22 | Scenario 23 | Scenario 24 | Scenario 25 | Scenario 26 | Scenario 27 | Scenario 28 | Scenario 29 | Scenario 30 | Scenario 31 | Scenario 32 | Scenario 33 | Scenario 34 | Scenario 35 | Scenario 36 | Scenario 37 | Scenario 38 | Scenario 39 | Scenario 40 | Scenario 41 | Scenario 42 | Scenario 43 | Scenario 44 | Scenario 45 | Scenario 46 | Scenario 47 | Scenario 48 | Scenario 49 | Scenario 50 | Scenario 51 | Scenario 52 | Scenario 53 | Scenario 54 | Scenario 55 | Scenario 56 | Scenario 57 | Scenario 58 | Scenario 59 | Scenario 60 | Scenario 61 | Scenario 62 | Scenario 63 | Scenario 64 | Scenario 65 | Scenario 66 | Scenario 67 | Scenario 68 | Scenario 69 | Scenario 70 | Scenario 71 | Scenario 72 | Scenario 73 | Scenario 74 | Scenario 75 | Scenario 76 | Scenario 77 | Scenario 78 | Scenario 79 | Scenario 80 | Scenario 81 | Scenario 82 | Scenario 83 | Scenario 84 | Scenario 85 | Scenario 86 | Scenario 87 | Scenario 88 | Scenario 89 | Scenario 90 | Scenario 91 | Scenario 92 | Scenario 93 | Scenario 94 | Scenario 95 | Scenario 96 | Scenario 97 | Scenario 98 | Scenario 99 | Scenario 100 | Scenario 101 | Scenario 102 | Scenario 103 | Scenario 104 | Scenario 105 | Scenario 106 | Scenario 107 | Scenario 108 | Scenario 109 | Scenario 110 | Scenario 111 | Scenario 112 | Scenario 113 | Scenario 114 | Scenario 115 | Scenario 116 | Scenario 117 | Scenario 118 | Scenario 119 | Scenario 120 | Scenario 121 | Scenario 122 | Scenario 123 | Scenario 124 | Scenario 125 | Scenario 126 | Scenario 127 | Scenario 128 | Scenario 129 | Scenario 130 | Scenario 131 | Scenario 132 | Scenario 133 | Scenario 134 | Scenario 135 | Scenario 136 | Scenario 137 | Scenario 138 | Scenario 139 | Scenario 140 | Scenario 141 | Scenario 142 | Scenario 143 | Scenario 144 | Scenario 145 | Scenario 146 | Scenario 147 | Scenario 148 | Scenario 149 | Scenario 150 | Scenario 151 | Scenario 152 | Scenario 153 | Scenario 154 | Scenario 155 | Scenario 156 | Scenario 157 | Scenario 158 | Scenario 159 | Scenario 160 | Scenario 161 | Scenario 162 | Scenario 163 | Scenario 164 | Scenario 165 | Scenario 166 | Scenario 167 | Scenario 168 | Scenario 169 | Scenario 170 | Scenario 171 | Scenario 172 | Scenario 173 | Scenario 174 | Scenario 175 | Scenario 176 | Scenario 177 | Scenario 178 | Scenario 179 | Scenario 180 | Scenario 181 | Scenario 182 | Scenario 183 | Scenario 184 | Scenario 185 | Scenario 186 | Scenario 187 | Scenario 188 | Scenario 189 | Scenario 190 | Scenario 191 | Scenario 192 | Scenario 193 | Scenario 194 | Scenario 195 | Scenario 196 | Scenario 197 | Scenario 198 | Scenario 199 | Scenario 200 | Scenario 201 | Scenario 202 | Scenario 203 | Scenario 204 | Scenario 205 | Scenario 206 | Scenario 207 | Scenario 208 | Scenario 209 | Scenario 210 | Scenario 211 | Scenario 212 | Scenario 213 | Scenario 214 | Scenario 215 | Scenario 216 | Scenario 217 | Scenario 218 | Scenario 219 | Scenario 220 | Scenario 221 | Scenario 222 | Scenario 223 | Scenario 224 | Scenario 225 | Scenario 226 | Scenario 227 | Scenario 228 | Scenario 229 | Scenario 230 | Scenario 231 | Scenario 232 | Scenario 233 | Scenario 234 | Scenario 235 | Scenario 236 | Scenario 237 | Scenario 238 | Scenario 239 | Scenario 240 | Scenario 241 | Scenario 242 | Scenario 243 | Scenario 244 | Scenario 245 | Scenario 246 | Scenario 247 | Scenario 248 | Scenario 249 | Scenario 250 | Scenario 251 | Scenario 252 | Scenario 253 | Scenario 254 | Scenario 255 | Scenario 256 | Scenario 257 | Scenario 258 | Scenario 259 | Scenario 260 | Scenario 261 | Scenario 262 | Scenario 263 | Scenario 264 | Scenario 265 | Scenario 266 | Scenario 267 | Scenario 268 | Scenario 269 | Scenario 270 | Scenario 271 | Scenario 272 | Scenario 273 | Scenario 274 | Scenario 275 | Scenario 276 | Scenario 277 | Scenario 278 | Scenario 279 | Scenario 280 | Scenario 281 | Scenario 282 | Scenario 283 | Scenario 284 | Scenario 285 | Scenario 286 | Scenario 287 | Scenario 288 | Scenario 289 | Scenario 290 | Scenario 291 | Scenario 292 | Scenario 293 | Scenario 294 | Scenario 295 | Scenario 296 | Scenario 297 | Scenario 298 | Scenario 299 | Scenario 300 | Scenario 301 | Scenario 302 | Scenario 303 | Scenario 304 | Scenario 305 | Scenario 306 | Scenario 307 | Scenario 308 | Scenario 309 | Scenario 310 | Scenario 311 | Scenario 312 | Scenario 313 | Scenario 314 | Scenario 315 | Scenario 316 | Scenario 317 | Scenario 318 | Scenario 319 | Scenario 320 | Scenario 321 | Scenario 322 | Scenario 323 | Scenario 324 | Scenario 325 | Scenario 326 | Scenario 327 | Scenario 328 | Scenario 329 | Scenario 330 | Scenario 331 | Scenario 332 | Scenario 333 | Scenario 334 | Scenario 335 | Scenario 336 | Scenario 337 | Scenario 338 | Scenario 339 | Scenario 340 | Scenario 341 | Scenario 342 | Scenario 343 | Scenario 344 | Scenario 345 | Scenario 346 | Scenario 347 | Scenario 348 | Scenario 349 | Scenario 350 | Scenario 351 | Scenario 352 | Scenario 353 | Scenario 354 | Scenario 355 | Scenario 356 | Scenario 357 | Scenario 358 | Scenario 359 | Scenario 360 | Scenario 361 | Scenario 362 | Scenario 363 | Scenario 364 | Scenario 365 | Scenario 366 | Scenario 367 | Scenario 368 | Scenario 369 | Scenario 370 | Scenario 371 | Scenario 372 | Scenario 373 | Scenario 374 | Scenario 375 | Scenario 376 | Scenario 377 | Scenario 378
[5 empty] | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated
Inner Ring Inner Dia |  | mm | 35.5 | 36.5 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44
Inner Ring Outer Dia |  | mm | 49 | 49 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65
Calcite X Normal | Minimize | N/mm^2 | 6.2e-05 | 9.3e-05 | 0.00011 | 0.000119 | 0.000123 | 0.000113 | 0.000126 | 0.000122 | 0.000124 | 0.00014 | 0.00019 | 0.000154 | 0.000126 | 0.000124 | 0.000133 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000113 | 9.8e-05 | 0.000105 | 0.000109 | 0.000103 | 0.00016 | 0.000261 | 0.000242 | 0.000264 | 0.000281 | 0.000148 | 0.000147 | 0.00023 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000113 | 0.000101 | 0.00013 | 0.000122 | 0.00012 | 0.00013 | 0.000114 | 0.000115 | 0.000121 | 0.000122 | 0.000118 | 0.000156 | 0.000102 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000235 | 0.000224 | 0.000234 | 0.000178 | 0.000241 | 0.000244 | 0.000239 | 0.000255 | 0.0001 | 0.000122 | 0.000131 | 0.000107 | 0.00024 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.2e-05 | 8.5e-05 | 9.3e-05 | 7.7e-05 | 8.3e-05 | 9.1e-05 | 0.000141 | 0.000155 | 7.5e-05 | 9.8e-05 | 8.4e-05 | 9.9e-05 | 0.000148 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000178 | 0.000171 | 0.000182 | 0.000224 | 0.00025 | 0.000223 | 0.00011 | 0.000114 | 0.000101 | 9.9e-05 | 0.000107 | 0.000172 | 0.000108 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000127 | 8.5e-05 | 0.000125 | 0.000115 | 0.000103 | 0.000102 | 0.000172 | 0.000114 | 0.000199 | 0.00024 | 0.000258 | 0.000261 | 0.000154 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000148 | 0.000172 | 0.00017 | 0.000163 | 0.000125 | 0.00019 | 0.00014 | 0.000181 | 0.000131 | 0.000143 | 0.000118 | 0.000134 | 0.000151 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.4e-05 | 0.000122 | 0.00012 | 0.000104 | 0.000103 | 0.000133 | 0.000114 | 0.000108 | 0.00012 | 0.000119 | 0.00015 | 0.00026 | 0.000223 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.9e-05 | 8.7e-05 | 0.000101 | 8.7e-05 | 9.1e-05 | 9.3e-05 | 0.000104 | 9.8e-05 | 0.000142 | 0.00017 | 0.000151 | 0.000147 | 0.000172 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 9.6e-05 | 8.4e-05 | 9.7e-05 | 9.6e-05 | 0.000103 | 0.000101 | 8.8e-05 | 0.000115 | 9.3e-05 | 0.000107 | 0.000118 | 9.6e-05 | 0.000133 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000179 | 0.000113 | 0.000105 | 0.000117 | 0.000147 | 0.00015 | 0.00019 | 0.000103 | 0.000127 | 0.000149 | 0.000144 | 0.000155 | 0.000113 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000107 | 0.000112 | 0.000184 | 7.9e-05 | 0.00012 | 7.9e-05 | 0.000117 | 9.4e-05 | 9.8e-05 | 0.0001 | 0.000113 | 0.000134 | 0.00017 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000135 | 0.000141 | 0.000131 | 0.000152 | 0.000107 | 0.000118 | 0.000134 | 0.000121 | 0.000132 | 0.000121 | 0.00011 | 0.000165 | 0.000132 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000135 | 0.000104 | 0.000148 | 0.000163 | 9.9e-05 | 0.000187 | 0.000109 | 9.8e-05 | 0.000161 | 0.000136 | 0.000103 | 0.000109 | 0.000129 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000101 | 8.9e-05 | 0.000104 | 0.000109 | 8.6e-05 | 0.000114 | 0.0001 | 0.00018 | 0.000132 | 0.000115 | 0.000142 | 0.000116 | 0.000154 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 9.4e-05 | 0.000207 | 0.000183 | 0.000183 | 0.000184 | 0.000143 | 0.000225 | 0.000166 | 9.8e-05 | 0.000119 | 9e-05 | 0.000166 | 0.000101 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000156 | 0.000125 | 0.000169 | 0.000134 | 0.000177 | 0.000128 | 0.000134 | 0.00015 | 0.000126 | 0.000121 | 0.000142 | 0.000168 | 0.000161 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000122 | 0.000132 | 0.000115 | 0.00011 | 0.000176 | 9.6e-05 | 0.000186 | 0.000198 | 0.000162 | 0.000184 | 0.000169 | 0.000177 | 0.000203 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000106 | 0.000103 | 0.000159 | 0.000114 | 9.3e-05 | 9.7e-05 | 0.000102 | 0.000105 | 0.000106 | 0.000113 | 0.000109 | 0.000126 | 0.00014 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000118 | 0.000125 | 0.000135 | 0.00014 | 0.00014 | 0.000137 | 0.000159 | 0.000161 | 0.000174 | 0.000174 | 0.000149 | 0.000158 | 0.000113 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite Y Normal | Minimize | N/mm^2 | 7.6e-05 | 9.1e-05 | 0.000106 | 0.000116 | 0.000123 | 8.3e-05 | 0.00013 | 7.7e-05 | 0.000128 | 0.000101 | 0.000112 | 9.4e-05 | 0.000126 | 0.000115 | 0.000127 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.6e-05 | 7.3e-05 | 8.5e-05 | 7.5e-05 | 9.6e-05 | 0.00016 | 0.000178 | 0.000148 | 0.000187 | 0.000217 | 0.000113 | 0.000118 | 0.000187 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000111 | 0.000119 | 0.000131 | 0.000116 | 0.000113 | 0.000111 | 0.000113 | 0.000124 | 0.00011 | 0.00012 | 0.000133 | 9.2e-05 | 0.000121 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000124 | 0.000128 | 0.000115 | 0.000152 | 0.000109 | 0.000107 | 0.000107 | 0.000115 | 7.5e-05 | 8.9e-05 | 0.000173 | 0.000121 | 0.000132 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.6e-05 | 9.2e-05 | 9.1e-05 | 7e-05 | 7.1e-05 | 7.4e-05 | 0.00011 | 9e-05 | 0.000115 | 0.000122 | 0.000131 | 0.000106 | 0.000217 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000105 | 0.000106 | 0.000108 | 0.000123 | 9.4e-05 | 0.000134 | 8.4e-05 | 9.1e-05 | 8.7e-05 | 0.000111 | 9.9e-05 | 0.000127 | 0.00012 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000132 | 0.000105 | 0.000145 | 0.000105 | 0.00016 | 0.000159 | 0.000107 | 8.9e-05 | 0.000119 | 0.000214 | 0.000123 | 0.000241 | 0.000144 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000107 | 0.000108 | 0.000108 | 0.000107 | 9.4e-05 | 0.000105 | 0.000119 | 0.000107 | 0.000125 | 0.000127 | 0.00017 | 0.000127 | 0.000164 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00011 | 8.3e-05 | 7.4e-05 | 0.000129 | 8.6e-05 | 9.1e-05 | 9.5e-05 | 9.3e-05 | 0.000106 | 0.000116 | 0.000123 | 0.000133 | 0.000138 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.3e-05 | 0.000165 | 0.000177 | 0.00017 | 0.000175 | 0.000178 | 0.000173 | 0.000168 | 9.8e-05 | 0.000197 | 0.000134 | 0.000129 | 0.000131 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000159 | 0.000152 | 0.00011 | 9.9e-05 | 0.000101 | 0.000111 | 7.7e-05 | 0.000121 | 9.6e-05 | 9.8e-05 | 0.000121 | 0.00013 | 0.000108 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000104 | 0.000103 | 0.000142 | 0.000139 | 0.000119 | 0.000113 | 0.000117 | 0.000114 | 8.9e-05 | 9.6e-05 | 0.000105 | 0.000111 | 0.00015 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000121 | 0.000122 | 9.4e-05 | 0.000123 | 0.000124 | 0.000109 | 0.000138 | 7.9e-05 | 8.8e-05 | 9.2e-05 | 8.6e-05 | 0.000105 | 0.000132 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000113 | 0.000118 | 0.000119 | 0.000118 | 0.000116 | 0.000124 | 0.000118 | 0.000115 | 9.1e-05 | 7.1e-05 | 7.6e-05 | 9.8e-05 | 0.000134 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000102 | 0.0001 | 0.000106 | 0.000109 | 0.000112 | 0.000122 | 0.000114 | 0.000103 | 0.000119 | 9.5e-05 | 0.000107 | 0.000117 | 0.000155 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.9e-05 | 0.0001 | 9.3e-05 | 0.000121 | 0.000125 | 0.000106 | 0.000134 | 0.000101 | 0.000124 | 0.000118 | 0.000128 | 0.00017 | 0.0001 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 9.7e-05 | 0.000139 | 0.00016 | 0.000159 | 0.000149 | 0.000132 | 0.00018 | 0.000149 | 0.000126 | 9.7e-05 | 0.000129 | 0.000161 | 0.000113 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000111 | 0.000125 | 0.000117 | 0.000147 | 0.000112 | 0.000128 | 0.000117 | 0.000135 | 0.000115 | 0.000114 | 0.000134 | 0.000121 | 0.000146 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000133 | 0.000145 | 0.000128 | 0.000153 | 0.000122 | 0.000146 | 0.000173 | 0.000161 | 0.000135 | 0.000176 | 0.000137 | 0.000141 | 0.000111 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 9e-05 | 7.4e-05 | 9.2e-05 | 0.000104 | 9e-05 | 9.3e-05 | 0.000142 | 0.000132 | 0.000168 | 0.000184 | 0.000188 | 0.000203 | 0.000187 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000172 | 0.000174 | 0.000182 | 0.000176 | 0.000169 | 0.000165 | 0.000122 | 0.00012 | 0.000139 | 0.000161 | 0.000104 | 0.000102 | 0.000159 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite XY Shear | Minimize | N/mm^2 | 4.5e-05 | 4.8e-05 | 5.8e-05 | 7.1e-05 | 8.2e-05 | 4.3e-05 | 7.6e-05 | 5.3e-05 | 7e-05 | 4.5e-05 | 5.8e-05 | 4.7e-05 | 8.3e-05 | 7.2e-05 | 7.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.1e-05 | 5.9e-05 | 6.2e-05 | 6.6e-05 | 5.5e-05 | 9.1e-05 | 7.8e-05 | 8.5e-05 | 7.3e-05 | 6.3e-05 | 4.8e-05 | 8.8e-05 | 9.4e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 5.9e-05 | 5.1e-05 | 7.4e-05 | 5.2e-05 | 7.9e-05 | 8.2e-05 | 8.1e-05 | 8.2e-05 | 7.8e-05 | 8.5e-05 | 8.1e-05 | 5.9e-05 | 6e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6e-05 | 6.4e-05 | 6.4e-05 | 6.9e-05 | 5.7e-05 | 5.7e-05 | 5.9e-05 | 5.9e-05 | 4.3e-05 | 4.9e-05 | 6.4e-05 | 8.5e-05 | 6e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 4.5e-05 | 6.7e-05 | 4.8e-05 | 4.3e-05 | 4.5e-05 | 4.7e-05 | 6.8e-05 | 6e-05 | 6e-05 | 5.3e-05 | 5.6e-05 | 5.7e-05 | 7.4e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 5e-05 | 5.3e-05 | 5.7e-05 | 5.8e-05 | 5.7e-05 | 5.4e-05 | 5.2e-05 | 5.8e-05 | 7.1e-05 | 8.6e-05 | 6.4e-05 | 6.6e-05 | 7.9e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.5e-05 | 5.4e-05 | 6.6e-05 | 5.9e-05 | 7.1e-05 | 7.3e-05 | 5.9e-05 | 4.4e-05 | 9.8e-05 | 6.9e-05 | 9.9e-05 | 6.2e-05 | 7e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.2e-05 | 7.8e-05 | 7.4e-05 | 6.1e-05 | 5.3e-05 | 7.6e-05 | 6.8e-05 | 6.7e-05 | 6.2e-05 | 6.6e-05 | 5.5e-05 | 5.7e-05 | 5.5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 5.6e-05 | 5.4e-05 | 5.3e-05 | 4.7e-05 | 5.3e-05 | 5.6e-05 | 5.2e-05 | 4.9e-05 | 4.4e-05 | 5.5e-05 | 5.4e-05 | 0.000113 | 7.5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.8e-05 | 4.3e-05 | 5.4e-05 | 4.1e-05 | 5.2e-05 | 5e-05 | 5.3e-05 | 4.8e-05 | 5.9e-05 | 5.2e-05 | 7.3e-05 | 6.2e-05 | 6.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 4.4e-05 | 4.4e-05 | 5.1e-05 | 5.1e-05 | 4.4e-05 | 4.7e-05 | 5.1e-05 | 4.4e-05 | 5.1e-05 | 4.9e-05 | 5.3e-05 | 5.3e-05 | 5.5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.8e-05 | 4.2e-05 | 5.4e-05 | 5.2e-05 | 8.6e-05 | 8.3e-05 | 6.7e-05 | 8.7e-05 | 6.2e-05 | 6.1e-05 | 6.2e-05 | 6.9e-05 | 5.4e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.3e-05 | 7.5e-05 | 7.5e-05 | 4.5e-05 | 6.1e-05 | 4.9e-05 | 5.7e-05 | 8.9e-05 | 8.2e-05 | 8.2e-05 | 6.6e-05 | 6.5e-05 | 9.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6e-05 | 6.3e-05 | 7.3e-05 | 6e-05 | 6.3e-05 | 7.1e-05 | 8.2e-05 | 8.5e-05 | 7.3e-05 | 9.5e-05 | 8.8e-05 | 7.4e-05 | 7.7e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 5.6e-05 | 5.5e-05 | 5.9e-05 | 6.3e-05 | 5.4e-05 | 8.4e-05 | 5.5e-05 | 7.3e-05 | 7.8e-05 | 7.9e-05 | 9.3e-05 | 7.1e-05 | 4.9e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 5.4e-05 | 5.8e-05 | 5.1e-05 | 7.6e-05 | 9e-05 | 5.6e-05 | 7.9e-05 | 4.8e-05 | 4.8e-05 | 4.9e-05 | 5e-05 | 5.5e-05 | 5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 4.7e-05 | 8.2e-05 | 8.8e-05 | 8.9e-05 | 8.5e-05 | 7.5e-05 | 6.6e-05 | 7.2e-05 | 6.9e-05 | 0.000102 | 6.2e-05 | 6.3e-05 | 5.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.6e-05 | 5.8e-05 | 6.7e-05 | 7.8e-05 | 7.1e-05 | 8.1e-05 | 6.9e-05 | 7.2e-05 | 7.5e-05 | 6.5e-05 | 7.5e-05 | 6.3e-05 | 7.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.2e-05 | 7.1e-05 | 6.8e-05 | 5.5e-05 | 6.5e-05 | 8.7e-05 | 6.4e-05 | 6.7e-05 | 6.4e-05 | 7.3e-05 | 7.4e-05 | 8.3e-05 | 5.5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.5e-05 | 5.1e-05 | 6.9e-05 | 8.1e-05 | 7.2e-05 | 6.7e-05 | 8.9e-05 | 9.1e-05 | 0.000101 | 0.000107 | 9.8e-05 | 0.000108 | 7.7e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 5.7e-05 | 5.1e-05 | 5.9e-05 | 5.9e-05 | 5.4e-05 | 5.1e-05 | 8.7e-05 | 8.5e-05 | 9.1e-05 | 5.2e-05 | 8.8e-05 | 4.5e-05 | 8.8e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite YZ Shear | Minimize | N/mm^2 | 8.4e-05 | 7.2e-05 | 9.2e-05 | 8.7e-05 | 8e-05 | 6.4e-05 | 7.7e-05 | 7.1e-05 | 7.8e-05 | 7.2e-05 | 6.5e-05 | 6.5e-05 | 5.9e-05 | 6.9e-05 | 7.2e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.9e-05 | 8.3e-05 | 8.8e-05 | 9e-05 | 7e-05 | 0.000179 | 0.000111 | 0.000113 | 0.000111 | 0.000129 | 7.5e-05 | 0.000105 | 9.7e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.6e-05 | 6.6e-05 | 0.000101 | 7e-05 | 7e-05 | 7.8e-05 | 7.1e-05 | 7.6e-05 | 6.6e-05 | 6.7e-05 | 8.5e-05 | 5.2e-05 | 6.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.6e-05 | 7.7e-05 | 7.5e-05 | 6.9e-05 | 7.3e-05 | 7.7e-05 | 7.9e-05 | 7.9e-05 | 7.7e-05 | 8e-05 | 0.000104 | 8.7e-05 | 0.000102 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.4e-05 | 8.4e-05 | 7.2e-05 | 8.7e-05 | 8.3e-05 | 8.1e-05 | 5.9e-05 | 9.8e-05 | 7.2e-05 | 5.6e-05 | 7e-05 | 6.4e-05 | 0.000112 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000104 | 0.000102 | 0.000105 | 8e-05 | 7.3e-05 | 7.9e-05 | 8.1e-05 | 8.6e-05 | 7.6e-05 | 7.8e-05 | 6e-05 | 6.9e-05 | 6.8e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7e-05 | 9.6e-05 | 7.6e-05 | 7.3e-05 | 7.4e-05 | 8.1e-05 | 9.4e-05 | 6.2e-05 | 0.000111 | 0.000101 | 0.000103 | 0.000107 | 9.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.8e-05 | 9.1e-05 | 9.4e-05 | 9.3e-05 | 5.6e-05 | 9.1e-05 | 9.1e-05 | 8.9e-05 | 9e-05 | 8.2e-05 | 0.000104 | 0.000109 | 0.0001 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.7e-05 | 6.6e-05 | 6.7e-05 | 0.000114 | 6.1e-05 | 6.5e-05 | 6.2e-05 | 6.7e-05 | 6.5e-05 | 5.8e-05 | 8.4e-05 | 9.3e-05 | 9.5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.2e-05 | 8.5e-05 | 8.1e-05 | 8.3e-05 | 8.4e-05 | 8e-05 | 8.6e-05 | 8.7e-05 | 9.1e-05 | 8.7e-05 | 8.7e-05 | 6.3e-05 | 9e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 6.7e-05 | 6.2e-05 | 9.2e-05 | 9.1e-05 | 9e-05 | 0.000102 | 0.00012 | 0.000109 | 0.000106 | 9.9e-05 | 0.00011 | 0.000105 | 0.000105 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 9.5e-05 | 5.5e-05 | 0.000126 | 0.000138 | 6.4e-05 | 6.4e-05 | 9.6e-05 | 6.9e-05 | 9.2e-05 | 9.2e-05 | 9e-05 | 7.1e-05 | 9.6e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 9.6e-05 | 9.3e-05 | 7.3e-05 | 9e-05 | 6.9e-05 | 8.1e-05 | 9.2e-05 | 7.6e-05 | 7.7e-05 | 7.7e-05 | 7.4e-05 | 7.5e-05 | 0.000102 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000124 | 0.000115 | 0.000112 | 0.000128 | 0.00012 | 0.000116 | 0.00012 | 0.000109 | 8.5e-05 | 9e-05 | 9.2e-05 | 9.1e-05 | 0.000133 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.2e-05 | 6e-05 | 8.3e-05 | 8.1e-05 | 9.4e-05 | 6.6e-05 | 8.8e-05 | 9.7e-05 | 6.9e-05 | 9.9e-05 | 0.00011 | 8.3e-05 | 9.8e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000107 | 8.2e-05 | 8.7e-05 | 9.6e-05 | 0.000145 | 7.6e-05 | 0.000106 | 6.7e-05 | 6.7e-05 | 6.5e-05 | 6.1e-05 | 7.4e-05 | 9.3e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.6e-05 | 9.6e-05 | 9.8e-05 | 9.8e-05 | 9.8e-05 | 8.9e-05 | 6.9e-05 | 8.9e-05 | 6.3e-05 | 6.6e-05 | 6.2e-05 | 6.3e-05 | 0.000112 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.6e-05 | 9e-05 | 7.4e-05 | 9.1e-05 | 7.6e-05 | 8.3e-05 | 7.4e-05 | 8.3e-05 | 8.3e-05 | 8.9e-05 | 0.000106 | 7.6e-05 | 0.000103 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.1e-05 | 9.2e-05 | 8.5e-05 | 0.000106 | 6.6e-05 | 0.00013 | 9.8e-05 | 0.000103 | 9.5e-05 | 0.000106 | 9.3e-05 | 0.000107 | 0.000105 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 8.7e-05 | 8.3e-05 | 9.3e-05 | 0.000102 | 8.4e-05 | 9.1e-05 | 0.000113 | 0.000107 | 0.000106 | 9.5e-05 | 0.000105 | 9.4e-05 | 8.5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 7.4e-05 | 7.3e-05 | 8.1e-05 | 7.3e-05 | 8.3e-05 | 7.9e-05 | 8.7e-05 | 8.6e-05 | 8.9e-05 | 0.000106 | 7.8e-05 | 0.000125 | 7.5e-05 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Displacement1 | Minimize | mm | 0.003694 | 0.00341 | 0.003856 | 0.003863 | 0.003847 | 0.003598 | 0.003894 | 0.003828 | 0.003973 | 0.003626 | 0.003765 | 0.003727 | 0.003924 | 0.003687 | 0.004087 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003794 | 0.003881 | 0.003859 | 0.003793 | 0.003864 | 0.003863 | 0.003893 | 0.003891 | 0.003746 | 0.0039 | 0.003916 | 0.003847 | 0.003923 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003858 | 0.003826 | 0.003572 | 0.004086 | 0.0042 | 0.003972 | 0.003928 | 0.003744 | 0.004023 | 0.003828 | 0.003878 | 0.004011 | 0.003871 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.004106 | 0.003834 | 0.003722 | 0.003739 | 0.004006 | 0.003921 | 0.00391 | 0.003849 | 0.003921 | 0.004034 | 0.0038 | 0.003937 | 0.003886 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003694 | 0.003488 | 0.00341 | 0.003688 | 0.003751 | 0.003699 | 0.003754 | 0.003463 | 0.003841 | 0.003831 | 0.003802 | 0.003673 | 0.003691 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00362 | 0.003629 | 0.00355 | 0.004011 | 0.00394 | 0.003907 | 0.003857 | 0.003942 | 0.004118 | 0.003926 | 0.003622 | 0.003899 | 0.003956 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003749 | 0.003823 | 0.003673 | 0.00392 | 0.003695 | 0.003394 | 0.003671 | 0.00366 | 0.00376 | 0.003775 | 0.00374 | 0.003788 | 0.003883 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003595 | 0.003498 | 0.003684 | 0.003624 | 0.003982 | 0.003592 | 0.003922 | 0.003557 | 0.003728 | 0.003912 | 0.0037 | 0.003736 | 0.003708 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003607 | 0.003946 | 0.003869 | 0.003884 | 0.003971 | 0.003825 | 0.004049 | 0.004001 | 0.003637 | 0.00383 | 0.003524 | 0.003768 | 0.003801 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003526 | 0.003773 | 0.003558 | 0.003798 | 0.003779 | 0.003771 | 0.003758 | 0.003577 | 0.003738 | 0.003746 | 0.00387 | 0.003816 | 0.003828 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003687 | 0.003749 | 0.003963 | 0.004 | 0.004087 | 0.003962 | 0.003738 | 0.004059 | 0.003633 | 0.004042 | 0.003555 | 0.003885 | 0.003709 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003666 | 0.003573 | 0.003876 | 0.003781 | 0.003984 | 0.003997 | 0.00381 | 0.00371 | 0.003643 | 0.003771 | 0.003792 | 0.00387 | 0.003758 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003935 | 0.003957 | 0.00365 | 0.003748 | 0.003808 | 0.003643 | 0.003896 | 0.00367 | 0.003818 | 0.003682 | 0.003664 | 0.00365 | 0.003721 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003886 | 0.003707 | 0.003898 | 0.003849 | 0.003652 | 0.003859 | 0.003855 | 0.003787 | 0.003762 | 0.004048 | 0.00366 | 0.003742 | 0.003654 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003878 | 0.003638 | 0.003604 | 0.003774 | 0.003844 | 0.003553 | 0.003757 | 0.003795 | 0.003613 | 0.003523 | 0.003628 | 0.003678 | 0.003701 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003779 | 0.003653 | 0.003547 | 0.003573 | 0.003932 | 0.003746 | 0.003949 | 0.003588 | 0.004132 | 0.004106 | 0.00418 | 0.004164 | 0.003719 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003585 | 0.003903 | 0.00392 | 0.003991 | 0.003938 | 0.003944 | 0.003587 | 0.004051 | 0.003871 | 0.003964 | 0.003879 | 0.003887 | 0.00373 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003643 | 0.003573 | 0.00396 | 0.003967 | 0.003725 | 0.003999 | 0.003925 | 0.003605 | 0.004063 | 0.003945 | 0.004033 | 0.003679 | 0.003908 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003882 | 0.003666 | 0.00377 | 0.003883 | 0.004042 | 0.003919 | 0.003801 | 0.003825 | 0.003857 | 0.00395 | 0.0039 | 0.004024 | 0.003777 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003992 | 0.003667 | 0.003868 | 0.003942 | 0.00399 | 0.003957 | 0.003864 | 0.003784 | 0.00393 | 0.003805 | 0.003896 | 0.003915 | 0.003939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003858 | 0.003922 | 0.003948 | 0.003515 | 0.004065 | 0.00385 | 0.003687 | 0.003764 | 0.003699 | 0.003453 | 0.003681 | 0.003397 | 0.003457 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error

## Calcite 10_8 optimization
Design Study 1
Scenarios/Iterations: | 1134 |  |  | optimal
Parameter Constraint or Goal | Format | Unit | Initial Value | Scenario 687 | Scenario 1 | Scenario 2 | Scenario 3 | Scenario 4 | Scenario 5 | Scenario 6 | Scenario 7 | Scenario 8 | Scenario 9 | Scenario 10 | Scenario 11 | Scenario 12 | Scenario 13 | Scenario 14 | Scenario 15 | Scenario 16 | Scenario 17 | Scenario 18 | Scenario 19 | Scenario 20 | Scenario 21 | Scenario 22 | Scenario 23 | Scenario 24 | Scenario 25 | Scenario 26 | Scenario 27 | Scenario 28 | Scenario 29 | Scenario 30 | Scenario 31 | Scenario 32 | Scenario 33 | Scenario 34 | Scenario 35 | Scenario 36 | Scenario 37 | Scenario 38 | Scenario 39 | Scenario 40 | Scenario 41 | Scenario 42 | Scenario 43 | Scenario 44 | Scenario 45 | Scenario 46 | Scenario 47 | Scenario 48 | Scenario 49 | Scenario 50 | Scenario 51 | Scenario 52 | Scenario 53 | Scenario 54 | Scenario 55 | Scenario 56 | Scenario 57 | Scenario 58 | Scenario 59 | Scenario 60 | Scenario 61 | Scenario 62 | Scenario 63 | Scenario 64 | Scenario 65 | Scenario 66 | Scenario 67 | Scenario 68 | Scenario 69 | Scenario 70 | Scenario 71 | Scenario 72 | Scenario 73 | Scenario 74 | Scenario 75 | Scenario 76 | Scenario 77 | Scenario 78 | Scenario 79 | Scenario 80 | Scenario 81 | Scenario 82 | Scenario 83 | Scenario 84 | Scenario 85 | Scenario 86 | Scenario 87 | Scenario 88 | Scenario 89 | Scenario 90 | Scenario 91 | Scenario 92 | Scenario 93 | Scenario 94 | Scenario 95 | Scenario 96 | Scenario 97 | Scenario 98 | Scenario 99 | Scenario 100 | Scenario 101 | Scenario 102 | Scenario 103 | Scenario 104 | Scenario 105 | Scenario 106 | Scenario 107 | Scenario 108 | Scenario 109 | Scenario 110 | Scenario 111 | Scenario 112 | Scenario 113 | Scenario 114 | Scenario 115 | Scenario 116 | Scenario 117 | Scenario 118 | Scenario 119 | Scenario 120 | Scenario 121 | Scenario 122 | Scenario 123 | Scenario 124 | Scenario 125 | Scenario 126 | Scenario 127 | Scenario 128 | Scenario 129 | Scenario 130 | Scenario 131 | Scenario 132 | Scenario 133 | Scenario 134 | Scenario 135 | Scenario 136 | Scenario 137 | Scenario 138 | Scenario 139 | Scenario 140 | Scenario 141 | Scenario 142 | Scenario 143 | Scenario 144 | Scenario 145 | Scenario 146 | Scenario 147 | Scenario 148 | Scenario 149 | Scenario 150 | Scenario 151 | Scenario 152 | Scenario 153 | Scenario 154 | Scenario 155 | Scenario 156 | Scenario 157 | Scenario 158 | Scenario 159 | Scenario 160 | Scenario 161 | Scenario 162 | Scenario 163 | Scenario 164 | Scenario 165 | Scenario 166 | Scenario 167 | Scenario 168 | Scenario 169 | Scenario 170 | Scenario 171 | Scenario 172 | Scenario 173 | Scenario 174 | Scenario 175 | Scenario 176 | Scenario 177 | Scenario 178 | Scenario 179 | Scenario 180 | Scenario 181 | Scenario 182 | Scenario 183 | Scenario 184 | Scenario 185 | Scenario 186 | Scenario 187 | Scenario 188 | Scenario 189 | Scenario 190 | Scenario 191 | Scenario 192 | Scenario 193 | Scenario 194 | Scenario 195 | Scenario 196 | Scenario 197 | Scenario 198 | Scenario 199 | Scenario 200 | Scenario 201 | Scenario 202 | Scenario 203 | Scenario 204 | Scenario 205 | Scenario 206 | Scenario 207 | Scenario 208 | Scenario 209 | Scenario 210 | Scenario 211 | Scenario 212 | Scenario 213 | Scenario 214 | Scenario 215 | Scenario 216 | Scenario 217 | Scenario 218 | Scenario 219 | Scenario 220 | Scenario 221 | Scenario 222 | Scenario 223 | Scenario 224 | Scenario 225 | Scenario 226 | Scenario 227 | Scenario 228 | Scenario 229 | Scenario 230 | Scenario 231 | Scenario 232 | Scenario 233 | Scenario 234 | Scenario 235 | Scenario 236 | Scenario 237 | Scenario 238 | Scenario 239 | Scenario 240 | Scenario 241 | Scenario 242 | Scenario 243 | Scenario 244 | Scenario 245 | Scenario 246 | Scenario 247 | Scenario 248 | Scenario 249 | Scenario 250 | Scenario 251 | Scenario 252 | Scenario 253 | Scenario 254 | Scenario 255 | Scenario 256 | Scenario 257 | Scenario 258 | Scenario 259 | Scenario 260 | Scenario 261 | Scenario 262 | Scenario 263 | Scenario 264 | Scenario 265 | Scenario 266 | Scenario 267 | Scenario 268 | Scenario 269 | Scenario 270 | Scenario 271 | Scenario 272 | Scenario 273 | Scenario 274 | Scenario 275 | Scenario 276 | Scenario 277 | Scenario 278 | Scenario 279 | Scenario 280 | Scenario 281 | Scenario 282 | Scenario 283 | Scenario 284 | Scenario 285 | Scenario 286 | Scenario 287 | Scenario 288 | Scenario 289 | Scenario 290 | Scenario 291 | Scenario 292 | Scenario 293 | Scenario 294 | Scenario 295 | Scenario 296 | Scenario 297 | Scenario 298 | Scenario 299 | Scenario 300 | Scenario 301 | Scenario 302 | Scenario 303 | Scenario 304 | Scenario 305 | Scenario 306 | Scenario 307 | Scenario 308 | Scenario 309 | Scenario 310 | Scenario 311 | Scenario 312 | Scenario 313 | Scenario 314 | Scenario 315 | Scenario 316 | Scenario 317 | Scenario 318 | Scenario 319 | Scenario 320 | Scenario 321 | Scenario 322 | Scenario 323 | Scenario 324 | Scenario 325 | Scenario 326 | Scenario 327 | Scenario 328 | Scenario 329 | Scenario 330 | Scenario 331 | Scenario 332 | Scenario 333 | Scenario 334 | Scenario 335 | Scenario 336 | Scenario 337 | Scenario 338 | Scenario 339 | Scenario 340 | Scenario 341 | Scenario 342 | Scenario 343 | Scenario 344 | Scenario 345 | Scenario 346 | Scenario 347 | Scenario 348 | Scenario 349 | Scenario 350 | Scenario 351 | Scenario 352 | Scenario 353 | Scenario 354 | Scenario 355 | Scenario 356 | Scenario 357 | Scenario 358 | Scenario 359 | Scenario 360 | Scenario 361 | Scenario 362 | Scenario 363 | Scenario 364 | Scenario 365 | Scenario 366 | Scenario 367 | Scenario 368 | Scenario 369 | Scenario 370 | Scenario 371 | Scenario 372 | Scenario 373 | Scenario 374 | Scenario 375 | Scenario 376 | Scenario 377 | Scenario 378 | Scenario 379 | Scenario 380 | Scenario 381 | Scenario 382 | Scenario 383 | Scenario 384 | Scenario 385 | Scenario 386 | Scenario 387 | Scenario 388 | Scenario 389 | Scenario 390 | Scenario 391 | Scenario 392 | Scenario 393 | Scenario 394 | Scenario 395 | Scenario 396 | Scenario 397 | Scenario 398 | Scenario 399 | Scenario 400 | Scenario 401 | Scenario 402 | Scenario 403 | Scenario 404 | Scenario 405 | Scenario 406 | Scenario 407 | Scenario 408 | Scenario 409 | Scenario 410 | Scenario 411 | Scenario 412 | Scenario 413 | Scenario 414 | Scenario 415 | Scenario 416 | Scenario 417 | Scenario 418 | Scenario 419 | Scenario 420 | Scenario 421 | Scenario 422 | Scenario 423 | Scenario 424 | Scenario 425 | Scenario 426 | Scenario 427 | Scenario 428 | Scenario 429 | Scenario 430 | Scenario 431 | Scenario 432 | Scenario 433 | Scenario 434 | Scenario 435 | Scenario 436 | Scenario 437 | Scenario 438 | Scenario 439 | Scenario 440 | Scenario 441 | Scenario 442 | Scenario 443 | Scenario 444 | Scenario 445 | Scenario 446 | Scenario 447 | Scenario 448 | Scenario 449 | Scenario 450 | Scenario 451 | Scenario 452 | Scenario 453 | Scenario 454 | Scenario 455 | Scenario 456 | Scenario 457 | Scenario 458 | Scenario 459 | Scenario 460 | Scenario 461 | Scenario 462 | Scenario 463 | Scenario 464 | Scenario 465 | Scenario 466 | Scenario 467 | Scenario 468 | Scenario 469 | Scenario 470 | Scenario 471 | Scenario 472 | Scenario 473 | Scenario 474 | Scenario 475 | Scenario 476 | Scenario 477 | Scenario 478 | Scenario 479 | Scenario 480 | Scenario 481 | Scenario 482 | Scenario 483 | Scenario 484 | Scenario 485 | Scenario 486 | Scenario 487 | Scenario 488 | Scenario 489 | Scenario 490 | Scenario 491 | Scenario 492 | Scenario 493 | Scenario 494 | Scenario 495 | Scenario 496 | Scenario 497 | Scenario 498 | Scenario 499 | Scenario 500 | Scenario 501 | Scenario 502 | Scenario 503 | Scenario 504 | Scenario 505 | Scenario 506 | Scenario 507 | Scenario 508 | Scenario 509 | Scenario 510 | Scenario 511 | Scenario 512 | Scenario 513 | Scenario 514 | Scenario 515 | Scenario 516 | Scenario 517 | Scenario 518 | Scenario 519 | Scenario 520 | Scenario 521 | Scenario 522 | Scenario 523 | Scenario 524 | Scenario 525 | Scenario 526 | Scenario 527 | Scenario 528 | Scenario 529 | Scenario 530 | Scenario 531 | Scenario 532 | Scenario 533 | Scenario 534 | Scenario 535 | Scenario 536 | Scenario 537 | Scenario 538 | Scenario 539 | Scenario 540 | Scenario 541 | Scenario 542 | Scenario 543 | Scenario 544 | Scenario 545 | Scenario 546 | Scenario 547 | Scenario 548 | Scenario 549 | Scenario 550 | Scenario 551 | Scenario 552 | Scenario 553 | Scenario 554 | Scenario 555 | Scenario 556 | Scenario 557 | Scenario 558 | Scenario 559 | Scenario 560 | Scenario 561 | Scenario 562 | Scenario 563 | Scenario 564 | Scenario 565 | Scenario 566 | Scenario 567 | Scenario 568 | Scenario 569 | Scenario 570 | Scenario 571 | Scenario 572 | Scenario 573 | Scenario 574 | Scenario 575 | Scenario 576 | Scenario 577 | Scenario 578 | Scenario 579 | Scenario 580 | Scenario 581 | Scenario 582 | Scenario 583 | Scenario 584 | Scenario 585 | Scenario 586 | Scenario 587 | Scenario 588 | Scenario 589 | Scenario 590 | Scenario 591 | Scenario 592 | Scenario 593 | Scenario 594 | Scenario 595 | Scenario 596 | Scenario 597 | Scenario 598 | Scenario 599 | Scenario 600 | Scenario 601 | Scenario 602 | Scenario 603 | Scenario 604 | Scenario 605 | Scenario 606 | Scenario 607 | Scenario 608 | Scenario 609 | Scenario 610 | Scenario 611 | Scenario 612 | Scenario 613 | Scenario 614 | Scenario 615 | Scenario 616 | Scenario 617 | Scenario 618 | Scenario 619 | Scenario 620 | Scenario 621 | Scenario 622 | Scenario 623 | Scenario 624 | Scenario 625 | Scenario 626 | Scenario 627 | Scenario 628 | Scenario 629 | Scenario 630 | Scenario 631 | Scenario 632 | Scenario 633 | Scenario 634 | Scenario 635 | Scenario 636 | Scenario 637 | Scenario 638 | Scenario 639 | Scenario 640 | Scenario 641 | Scenario 642 | Scenario 643 | Scenario 644 | Scenario 645 | Scenario 646 | Scenario 647 | Scenario 648 | Scenario 649 | Scenario 650 | Scenario 651 | Scenario 652 | Scenario 653 | Scenario 654 | Scenario 655 | Scenario 656 | Scenario 657 | Scenario 658 | Scenario 659 | Scenario 660 | Scenario 661 | Scenario 662 | Scenario 663 | Scenario 664 | Scenario 665 | Scenario 666 | Scenario 667 | Scenario 668 | Scenario 669 | Scenario 670 | Scenario 671 | Scenario 672 | Scenario 673 | Scenario 674 | Scenario 675 | Scenario 676 | Scenario 677 | Scenario 678 | Scenario 679 | Scenario 680 | Scenario 681 | Scenario 682 | Scenario 683 | Scenario 684 | Scenario 685 | Scenario 686 | Scenario 687 | Scenario 688 | Scenario 689 | Scenario 690 | Scenario 691 | Scenario 692 | Scenario 693 | Scenario 694 | Scenario 695 | Scenario 696 | Scenario 697 | Scenario 698 | Scenario 699 | Scenario 700 | Scenario 701 | Scenario 702 | Scenario 703 | Scenario 704 | Scenario 705 | Scenario 706 | Scenario 707 | Scenario 708 | Scenario 709 | Scenario 710 | Scenario 711 | Scenario 712 | Scenario 713 | Scenario 714 | Scenario 715 | Scenario 716 | Scenario 717 | Scenario 718 | Scenario 719 | Scenario 720 | Scenario 721 | Scenario 722 | Scenario 723 | Scenario 724 | Scenario 725 | Scenario 726 | Scenario 727 | Scenario 728 | Scenario 729 | Scenario 730 | Scenario 731 | Scenario 732 | Scenario 733 | Scenario 734 | Scenario 735 | Scenario 736 | Scenario 737 | Scenario 738 | Scenario 739 | Scenario 740 | Scenario 741 | Scenario 742 | Scenario 743 | Scenario 744 | Scenario 745 | Scenario 746 | Scenario 747 | Scenario 748 | Scenario 749 | Scenario 750 | Scenario 751 | Scenario 752 | Scenario 753 | Scenario 754 | Scenario 755 | Scenario 756 | Scenario 757 | Scenario 758 | Scenario 759 | Scenario 760 | Scenario 761 | Scenario 762 | Scenario 763 | Scenario 764 | Scenario 765 | Scenario 766 | Scenario 767 | Scenario 768 | Scenario 769 | Scenario 770 | Scenario 771 | Scenario 772 | Scenario 773 | Scenario 774 | Scenario 775 | Scenario 776 | Scenario 777 | Scenario 778 | Scenario 779 | Scenario 780 | Scenario 781 | Scenario 782 | Scenario 783 | Scenario 784 | Scenario 785 | Scenario 786 | Scenario 787 | Scenario 788 | Scenario 789 | Scenario 790 | Scenario 791 | Scenario 792 | Scenario 793 | Scenario 794 | Scenario 795 | Scenario 796 | Scenario 797 | Scenario 798 | Scenario 799 | Scenario 800 | Scenario 801 | Scenario 802 | Scenario 803 | Scenario 804 | Scenario 805 | Scenario 806 | Scenario 807 | Scenario 808 | Scenario 809 | Scenario 810 | Scenario 811 | Scenario 812 | Scenario 813 | Scenario 814 | Scenario 815 | Scenario 816 | Scenario 817 | Scenario 818 | Scenario 819 | Scenario 820 | Scenario 821 | Scenario 822 | Scenario 823 | Scenario 824 | Scenario 825 | Scenario 826 | Scenario 827 | Scenario 828 | Scenario 829 | Scenario 830 | Scenario 831 | Scenario 832 | Scenario 833 | Scenario 834 | Scenario 835 | Scenario 836 | Scenario 837 | Scenario 838 | Scenario 839 | Scenario 840 | Scenario 841 | Scenario 842 | Scenario 843 | Scenario 844 | Scenario 845 | Scenario 846 | Scenario 847 | Scenario 848 | Scenario 849 | Scenario 850 | Scenario 851 | Scenario 852 | Scenario 853 | Scenario 854 | Scenario 855 | Scenario 856 | Scenario 857 | Scenario 858 | Scenario 859 | Scenario 860 | Scenario 861 | Scenario 862 | Scenario 863 | Scenario 864 | Scenario 865 | Scenario 866 | Scenario 867 | Scenario 868 | Scenario 869 | Scenario 870 | Scenario 871 | Scenario 872 | Scenario 873 | Scenario 874 | Scenario 875 | Scenario 876 | Scenario 877 | Scenario 878 | Scenario 879 | Scenario 880 | Scenario 881 | Scenario 882 | Scenario 883 | Scenario 884 | Scenario 885 | Scenario 886 | Scenario 887 | Scenario 888 | Scenario 889 | Scenario 890 | Scenario 891 | Scenario 892 | Scenario 893 | Scenario 894 | Scenario 895 | Scenario 896 | Scenario 897 | Scenario 898 | Scenario 899 | Scenario 900 | Scenario 901 | Scenario 902 | Scenario 903 | Scenario 904 | Scenario 905 | Scenario 906 | Scenario 907 | Scenario 908 | Scenario 909 | Scenario 910 | Scenario 911 | Scenario 912 | Scenario 913 | Scenario 914 | Scenario 915 | Scenario 916 | Scenario 917 | Scenario 918 | Scenario 919 | Scenario 920 | Scenario 921 | Scenario 922 | Scenario 923 | Scenario 924 | Scenario 925 | Scenario 926 | Scenario 927 | Scenario 928 | Scenario 929 | Scenario 930 | Scenario 931 | Scenario 932 | Scenario 933 | Scenario 934 | Scenario 935 | Scenario 936 | Scenario 937 | Scenario 938 | Scenario 939 | Scenario 940 | Scenario 941 | Scenario 942 | Scenario 943 | Scenario 944 | Scenario 945 | Scenario 946 | Scenario 947 | Scenario 948 | Scenario 949 | Scenario 950 | Scenario 951 | Scenario 952 | Scenario 953 | Scenario 954 | Scenario 955 | Scenario 956 | Scenario 957 | Scenario 958 | Scenario 959 | Scenario 960 | Scenario 961 | Scenario 962 | Scenario 963 | Scenario 964 | Scenario 965 | Scenario 966 | Scenario 967 | Scenario 968 | Scenario 969 | Scenario 970 | Scenario 971 | Scenario 972 | Scenario 973 | Scenario 974 | Scenario 975 | Scenario 976 | Scenario 977 | Scenario 978 | Scenario 979 | Scenario 980 | Scenario 981 | Scenario 982 | Scenario 983 | Scenario 984 | Scenario 985 | Scenario 986 | Scenario 987 | Scenario 988 | Scenario 989 | Scenario 990 | Scenario 991 | Scenario 992 | Scenario 993 | Scenario 994 | Scenario 995 | Scenario 996 | Scenario 997 | Scenario 998 | Scenario 999 | Scenario 1000 | Scenario 1001 | Scenario 1002 | Scenario 1003 | Scenario 1004 | Scenario 1005 | Scenario 1006 | Scenario 1007 | Scenario 1008 | Scenario 1009 | Scenario 1010 | Scenario 1011 | Scenario 1012 | Scenario 1013 | Scenario 1014 | Scenario 1015 | Scenario 1016 | Scenario 1017 | Scenario 1018 | Scenario 1019 | Scenario 1020 | Scenario 1021 | Scenario 1022 | Scenario 1023 | Scenario 1024 | Scenario 1025 | Scenario 1026 | Scenario 1027 | Scenario 1028 | Scenario 1029 | Scenario 1030 | Scenario 1031 | Scenario 1032 | Scenario 1033 | Scenario 1034 | Scenario 1035 | Scenario 1036 | Scenario 1037 | Scenario 1038 | Scenario 1039 | Scenario 1040 | Scenario 1041 | Scenario 1042 | Scenario 1043 | Scenario 1044 | Scenario 1045 | Scenario 1046 | Scenario 1047 | Scenario 1048 | Scenario 1049 | Scenario 1050 | Scenario 1051 | Scenario 1052 | Scenario 1053 | Scenario 1054 | Scenario 1055 | Scenario 1056 | Scenario 1057 | Scenario 1058 | Scenario 1059 | Scenario 1060 | Scenario 1061 | Scenario 1062 | Scenario 1064 | Scenario 1065 | Scenario 1066 | Scenario 1067 | Scenario 1068 | Scenario 1069 | Scenario 1070 | Scenario 1071 | Scenario 1072 | Scenario 1076 | Scenario 1077 | Scenario 1078 | Scenario 1079 | Scenario 1080 | Scenario 1081 | Scenario 1082 | Scenario 1083 | Scenario 1084 | Scenario 1085 | Scenario 1086 | Scenario 1087 | Scenario 1088 | Scenario 1089 | Scenario 1090 | Scenario 1091 | Scenario 1092 | Scenario 1093 | Scenario 1094 | Scenario 1095 | Scenario 1096 | Scenario 1097 | Scenario 1098 | Scenario 1099 | Scenario 1100 | Scenario 1101 | Scenario 1102 | Scenario 1103 | Scenario 1104 | Scenario 1105 | Scenario 1106 | Scenario 1107 | Scenario 1108 | Scenario 1109 | Scenario 1110 | Scenario 1111 | Scenario 1112 | Scenario 1113 | Scenario 1114 | Scenario 1115 | Scenario 1116 | Scenario 1117 | Scenario 1118 | Scenario 1119 | Scenario 1120 | Scenario 1121 | Scenario 1122 | Scenario 1123 | Scenario 1124 | Scenario 1125 | Scenario 1126 | Scenario 1127 | Scenario 1128 | Scenario 1129 | Scenario 1130 | Scenario 1131 | Scenario 1132 | Scenario 1133 | Scenario 1134
|  |  |  | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Estimated | Estimated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Calculated | Estimated | Calculated | Calculated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Calculated | Estimated | Calculated | Estimated | Estimated | Calculated | Estimated | Estimated | Calculated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Calculated | Estimated | Calculated | Estimated | Calculated | Estimated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Estimated | Calculated | Calculated | Estimated | Calculated | Estimated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Estimated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Calculated | Estimated | Estimated | Estimated | Estimated | Estimated
Inner Ring Inner Dia |  | mm | 35.5 | 36.5 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44
Inner Ring Outer Dia |  | mm | 45 | 62 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 45 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 46 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 47 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 48 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 49 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 51 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 52 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 53 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 54 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 55 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 56 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 57 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 58 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 59 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 60 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 61 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 62 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 63 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 64 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65 | 65
Adhesive Patch D |  | mm | 4 | 4 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
Calcite X Normal | Minimize | N/mm^2 | 0.03168 | 0.000199 | 0.035309 | Simulation 0 to 50 C fails. | 0.035309 | 0.035309 | 0.035309 | 0.035315 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | Simulation 0 to 50 C fails. | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | Simulation 0 to 50 C fails. | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035314 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Simulation 0 to 50 C fails. | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.03531 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035315 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | Simulation 0 to 50 C fails. | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.03531 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035314 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | 0.035309 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.002208 | 0.010872 | 0.000353 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.002205 | 0.002205 | 0.00098 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035313 | 0.035314 | 0.035316 | 0.035313 | 0.035313 | 0.035313 | 0.035316 | 0.035313 | 0.035313 | 0.035314 | 0.035334 | 0.035318 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035313 | 0.035313 | 0.035313 | 0.035316 | 0.035313 | 0.035313 | 0.035313 | 0.035314 | 0.035313 | 0.035313 | 0.035313 | 0.035314 | 0.035314 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035313 | 0.035313 | 0.035314 | 0.035314 | 0.035314 | 0.035313 | 0.035313 | 0.035314 | 0.035314 | 0.035313 | 0.035314 | 0.035314 | 0.035314 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.031682 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.031683 | 0.03168 | 0.03168 | Simulation 0 to 50 C fails. | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.031683 | 0.031683 | 0.031683 | Simulation 0 to 50 C fails. | 0.031683 | 0.03168 | 0.000384 | 0.031683 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | Simulation 0 to 50 C fails. | 0.031683 | 0.03168 | 0.007917 | 0.031683 | Simulation 0 to 50 C fails. | 0.031683 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.031683 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.03168 | Simulation 0 to 50 C fails. | 0.007917 | Simulation 0 to 50 C fails. | 0.03168 | 0.007917 | Simulation 0 to 50 C fails. | 0.03168 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | 0.031682 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.031683 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | 0.031683 | 0.03168 | 0.031682 | 0.03168 | 0.03168 | 0.031682 | 0.031683 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.031683 | 0.031683 | 0.03168 | 0.03168 | 0.03168 | 0.031683 | 0.031683 | 0.03168 | 0.03168 | 0.031683 | 0.03168 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.031683 | Simulation 0 to 50 C fails. | 0.03168 | Simulation 0 to 50 C fails. | 0.031683 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.03168 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.007917 | Simulation 0 to 50 C fails. | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.007917 | Simulation 0 to 50 C fails. | 0.007917 | 0.007917 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.031683 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | Simulation 0 to 50 C fails. | 0.03168 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.03168 | 0.031683 | 0.03168 | 0.031683 | Simulation 0 to 50 C fails. | 0.03168 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.014533 | Simulation 0 to 50 C fails. | 0.000199 | 0.031684 | 0.000505 | 0.000627 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.031686 | 0.024095 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.031685 | 0.031685 | 0.031685 | 0.031698 | 0.031708 | 0.031709 | 0.031686 | 0.031694 | 0.031703 | 0.031684 | 0.031708 | 0.031684 | 0.031685 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.031685 | 0.031685 | 0.031684 | 0.031685 | 0.031696 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.031685 | 0.031685 | 0.031684 | 0.031685 | 0.031685 | 0.031685 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | 0.031684 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026008 | 0.026009 | 0.026009 | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026012 | 0.026009 | 0.026009 | 0.026009 | 0.026013 | 0.02601 | 0.026013 | 0.026011 | 0.026011 | 0.026011 | 0.026012 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026008 | 0.026012 | 0.026009 | 0.026009 | 0.026012 | 0.026011 | 0.026013 | 0.02601 | Simulation 0 to 50 C fails. | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.02601 | 0.026011 | 0.026008 | 0.026009 | Simulation 0 to 50 C fails. | 0.026012 | 0.02601 | 0.026013 | 0.026013 | 0.02601 | 0.026011 | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026008 | 0.026009 | 0.026012 | 0.026012 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026011 | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.026009 | Simulation 0 to 50 C fails. | 0.026009 | Simulation 0 to 50 C fails. | 0.026013 | Simulation 0 to 50 C fails. | 0.026013 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026011 | 0.026009 | 0.026009 | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026011 | 0.026014 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026008 | 0.026009 | Simulation 0 to 50 C fails. | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026011 | 0.026011 | 0.026008 | 0.026012 | 0.026009 | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026014 | Simulation 0 to 50 C fails. | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000315 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.006499 | 0.026012 | Simulation 0 to 50 C fails. | 0.013109 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.02601 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026008 | 0.026009 | 0.026009 | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026011 | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.026008 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.0065 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026011 | 0.026008 | 0.026009 | 0.026009 | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026011 | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026011 | 0.026008 | 0.026012 | 0.026012 | 0.026009 | 0.026009 | 0.02601 | 0.026013 | 0.02601 | 0.026011 | Simulation 0 to 50 C fails. | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026008 | 0.026009 | 0.026009 | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.026011 | 0.026011 | 0.026011 | 0.026011 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026008 | 0.026008 | 0.026008 | 0.026009 | 0.026012 | 0.026009 | 0.02601 | 0.02601 | 0.02601 | 0.026014 | 0.026011 | 0.026011 | 0.026012 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.026247 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026011 | 0.026012 | 0.026012 | 0.026012 | 0.026012 | 0.026013 | 0.026014 | 0.026013 | 0.026014 | 0.026014 | 0.026014 | 0.026016 | 0.026015 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026011 | 0.026011 | 0.026012 | 0.026012 | 0.026012 | 0.026012 | 0.026013 | 0.026013 | 0.026013 | 0.026014 | 0.026014 | 0.026014 | 0.026015 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.026011 | 0.026011 | 0.026012 | 0.026012 | 0.026012 | 0.026013 | 0.026013 | 0.026013 | 0.026013 | 0.026014 | 0.026014 | 0.026014 | 0.026015 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite Y Normal | Minimize | N/mm^2 | 0.020806 | 0.00013 | 0.022853 | Simulation 0 to 50 C fails. | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | Simulation 0 to 50 C fails. | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | Simulation 0 to 50 C fails. | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022855 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Simulation 0 to 50 C fails. | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022855 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | Simulation 0 to 50 C fails. | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022855 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | 0.022853 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.001429 | 0.007041 | 0.000229 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.001429 | 0.001429 | 0.000635 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | 0.022855 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020808 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020808 | 0.020806 | 0.020806 | Simulation 0 to 50 C fails. | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.020808 | 0.020808 | 0.020808 | Simulation 0 to 50 C fails. | 0.020808 | 0.020806 | 0.000251 | 0.020808 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | Simulation 0 to 50 C fails. | 0.020808 | 0.020806 | 0.005201 | 0.020808 | Simulation 0 to 50 C fails. | 0.020808 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020808 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.020806 | Simulation 0 to 50 C fails. | 0.005201 | Simulation 0 to 50 C fails. | 0.020806 | 0.005201 | Simulation 0 to 50 C fails. | 0.020806 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | 0.020808 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020808 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | 0.020808 | 0.020806 | 0.020808 | 0.020806 | 0.020806 | 0.020808 | 0.020808 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020808 | 0.020808 | 0.020806 | 0.020806 | 0.020806 | 0.020808 | 0.020808 | 0.020806 | 0.020806 | 0.020808 | 0.020806 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020808 | Simulation 0 to 50 C fails. | 0.020806 | Simulation 0 to 50 C fails. | 0.020808 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.020806 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.005201 | Simulation 0 to 50 C fails. | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.005201 | Simulation 0 to 50 C fails. | 0.005201 | 0.005201 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020808 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | Simulation 0 to 50 C fails. | 0.020806 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020806 | 0.020808 | 0.020806 | 0.020808 | Simulation 0 to 50 C fails. | 0.020806 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.009547 | Simulation 0 to 50 C fails. | 0.00013 | 0.020808 | 0.000331 | 0.000411 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.020808 | 0.015826 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | 0.020808 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019618 | 0.019616 | 0.019616 | 0.019616 | 0.019618 | 0.019615 | 0.019618 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019616 | 0.019618 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019618 | 0.019615 | Simulation 0 to 50 C fails. | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019619 | 0.019618 | 0.019616 | 0.019616 | Simulation 0 to 50 C fails. | 0.019618 | 0.019616 | 0.019618 | 0.019618 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019618 | 0.019618 | 0.019616 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.019616 | Simulation 0 to 50 C fails. | 0.019616 | Simulation 0 to 50 C fails. | 0.019618 | Simulation 0 to 50 C fails. | 0.019617 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019618 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019615 | 0.019615 | 0.019615 | 0.019617 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019616 | 0.019616 | Simulation 0 to 50 C fails. | 0.019616 | 0.019616 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019619 | 0.019618 | 0.019616 | 0.019618 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019615 | 0.019615 | 0.019617 | Simulation 0 to 50 C fails. | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00024 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.004911 | 0.019618 | Simulation 0 to 50 C fails. | 0.009894 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.019615 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.019616 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.004911 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019618 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019619 | 0.019616 | 0.019619 | 0.019618 | 0.019616 | 0.019616 | 0.019616 | 0.019618 | 0.019615 | 0.019615 | Simulation 0 to 50 C fails. | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019616 | 0.019616 | 0.019616 | 0.019616 | 0.019618 | 0.019616 | 0.019616 | 0.019616 | 0.019615 | 0.019617 | 0.019615 | 0.019615 | 0.019615 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.019618 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019617 | 0.019617 | 0.019617 | 0.019617 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019617 | 0.019617 | 0.019617 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019618 | 0.019617 | 0.019617 | 0.019617 | 0.019617 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite XY Shear | Minimize | psi | 3.421 | 0.021399 | 3.9877 | Simulation 0 to 50 C fails. | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | Simulation 0 to 50 C fails. | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | Simulation 0 to 50 C fails. | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9878 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Simulation 0 to 50 C fails. | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9878 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | Simulation 0 to 50 C fails. | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | 3.9877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.24912 | 1.2281 | 0.039839 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.24913 | 0.24913 | 0.1107 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9876 | 3.9878 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | 3.9878 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.421 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.4211 | 3.421 | 3.421 | Simulation 0 to 50 C fails. | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 3.4211 | 3.4211 | 3.421 | Simulation 0 to 50 C fails. | 3.421 | 3.421 | 0.041243 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | Simulation 0 to 50 C fails. | 3.421 | 3.421 | 0.85476 | 3.421 | Simulation 0 to 50 C fails. | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.421 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 3.421 | Simulation 0 to 50 C fails. | 0.85477 | Simulation 0 to 50 C fails. | 3.421 | 0.85476 | Simulation 0 to 50 C fails. | 3.421 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | 3.4211 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.421 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | 3.4211 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.4209 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.4211 | 3.4211 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.4211 | Simulation 0 to 50 C fails. | 3.421 | Simulation 0 to 50 C fails. | 3.421 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 3.421 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.85474 | Simulation 0 to 50 C fails. | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.4209 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.85479 | Simulation 0 to 50 C fails. | 0.85478 | 0.85477 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.421 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | Simulation 0 to 50 C fails. | 3.4209 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.4209 | 3.4209 | 3.4209 | 3.4209 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | Simulation 0 to 50 C fails. | 3.4209 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.5692 | Simulation 0 to 50 C fails. | 0.021399 | 3.4211 | 0.054358 | 0.067503 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 3.4211 | 2.6016 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.4212 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4212 | 3.4211 | 3.421 | 3.4211 | 3.421 | 3.421 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.4212 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.421 | 3.421 | 3.421 | 3.421 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 3.4212 | 3.4212 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.4211 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | 3.421 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8964 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8964 | 2.8963 | 2.8963 | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8964 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8964 | 2.8964 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.035026 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.72399 | 2.8963 | Simulation 0 to 50 C fails. | 1.4599 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 2.8963 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 2.8964 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.72399 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8964 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8964 | 2.8963 | 2.8964 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Simulation 0 to 50 C fails. | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | 2.8963 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 2.8964 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8965 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | 2.8964 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite YZ Shear | Minimize | N/mm^2 | 0.013939 | 8.8e-05 | 0.019079 | Simulation 0 to 50 C fails. | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | Simulation 0 to 50 C fails. | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | Simulation 0 to 50 C fails. | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.01908 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | Simulation 0 to 50 C fails. | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | Simulation 0 to 50 C fails. | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019078 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.00119 | 0.005874 | 0.000189 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.001189 | 0.001189 | 0.000528 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.019079 | 0.019079 | 0.019079 | 0.019078 | 0.019079 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019078 | 0.019077 | 0.019077 | 0.019077 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.01394 | 0.013939 | 0.013939 | Simulation 0 to 50 C fails. | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013938 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.01394 | 0.01394 | 0.01394 | Simulation 0 to 50 C fails. | 0.01394 | 0.013939 | 0.000169 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | Simulation 0 to 50 C fails. | 0.01394 | 0.013939 | 0.003485 | 0.013939 | Simulation 0 to 50 C fails. | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013938 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.013939 | Simulation 0 to 50 C fails. | 0.003485 | Simulation 0 to 50 C fails. | 0.013939 | 0.003485 | Simulation 0 to 50 C fails. | 0.013939 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | 0.01394 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | 0.01394 | 0.013939 | 0.01394 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013938 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.01394 | 0.01394 | 0.013939 | 0.013939 | 0.013939 | 0.01394 | 0.01394 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.01394 | Simulation 0 to 50 C fails. | 0.013939 | Simulation 0 to 50 C fails. | 0.01394 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.013939 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.003484 | Simulation 0 to 50 C fails. | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013938 | 0.013938 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.003485 | Simulation 0 to 50 C fails. | 0.003485 | 0.003485 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Simulation 0 to 50 C fails. | 0.013939 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | 0.013939 | Simulation 0 to 50 C fails. | 0.013939 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.006394 | Simulation 0 to 50 C fails. | 8.8e-05 | 0.013937 | 0.000222 | 0.000276 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.013938 | 0.010601 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013937 | 0.013937 | 0.013937 | 0.013937 | 0.013937 | 0.013937 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013937 | 0.013937 | 0.013937 | 0.013937 | 0.013937 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013937 | 0.013937 | 0.013937 | 0.013937 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | 0.013938 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | 0.013453 | Simulation 0 to 50 C fails. | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013454 | 0.013454 | 0.013453 | 0.013453 | Simulation 0 to 50 C fails. | 0.013454 | 0.013453 | 0.013454 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.013453 | Simulation 0 to 50 C fails. | 0.013453 | Simulation 0 to 50 C fails. | 0.013454 | Simulation 0 to 50 C fails. | 0.013454 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Simulation 0 to 50 C fails. | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013454 | 0.013454 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | Simulation 0 to 50 C fails. | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000163 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.003363 | 0.013454 | Simulation 0 to 50 C fails. | 0.006782 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.013453 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.013453 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.003363 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013454 | 0.013453 | 0.013454 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | Simulation 0 to 50 C fails. | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | 0.013453 | 0.013454 | 0.013453 | 0.013453 | 0.013453 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.013455 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | 0.013454 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite Z Disp | Minimize | mm | 0.002167 | 0.13283 | 0.001476 | Simulation 0 to 50 C fails. | 0.001077 | 0.001315 | 0.001671 | 0.002542 | 0.001196 | 0.001133 | 0.001592 | 0.001401 | 0.001288 | 0.001259 | 0.001765 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00059 | 0.001372 | Simulation 0 to 50 C fails. | 0.001136 | 0.000785 | 0.002168 | 0.001589 | 0.001169 | 0.000562 | 0.001629 | 0.001645 | 0.000946 | 0.001454 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.004781 | -0.001909 | Simulation 0 to 50 C fails. | 0.001356 | 0.001662 | 0.005242 | 3.53e-06 | 0.001579 | 0.001459 | 0.001553 | 0.001107 | 0.000703 | 0.0024 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001536 | 0.007754 | -0.000438 | 0.000732 | -0.000466 | 0.000551 | -4.6e-05 | 0.001364 | 0.001893 | 0.000339 | 0.000643 | 0.001854 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000986 | 0.001034 | -0.001462 | 0.004747 | 0.000971 | 0.001257 | -9.1e-05 | 0.000612 | 0.00259 | 0.001548 | Simulation 0 to 50 C fails. | -0.000541 | 0.00012 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002526 | 0.001866 | 0.001649 | 0.001501 | 0.00106 | 0.00148 | 0.00165 | 0.00162 | 0.001234 | 0.001199 | 0.001486 | -4.9e-05 | 0.004207 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000481 | 0.000564 | -0.002941 | -0.000379 | 0.000278 | 0.001575 | 0.000623 | 0.002098 | 0.000533 | 0.000637 | 0.002098 | 0.000436 | 0.000732 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001363 | 3.1e-05 | 0.001393 | 0.00141 | 0.001448 | 0.001363 | 0.001623 | 0.001666 | 0.00158 | 0.001288 | 0.001533 | 0.001598 | 0.001665 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001391 | 0.000966 | 0.001269 | 0.001437 | 0.001388 | 0.002414 | 0.001325 | 0.000877 | 0.001842 | -0.000973 | 0.001376 | 0.001752 | 0.001675 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000421 | -0.000844 | 0.001763 | 0.00164 | 0.00196 | 0.000506 | 0.001194 | 0.001269 | -0.000194 | -0.001671 | 0.003216 | 0.001375 | 0.001517 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.004322 | 0.001242 | Simulation 0 to 50 C fails. | -0.000139 | 0.000712 | 0.001287 | 0.00096 | 0.002409 | 0.001785 | 0.000327 | 0.005525 | 0.001743 | 0.001784 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00175 | 0.002082 | 0.00387 | 0.001514 | 0.001685 | 0.001817 | 0.000707 | 0.001407 | 0.00202 | 0.001019 | 0.002014 | 0.001191 | 0.001355 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001831 | 0.001684 | -0.001493 | 0.001507 | -0.000258 | 0.000454 | 0.001219 | 0.001372 | 0.000142 | 0.006925 | -0.000846 | 0.001046 | 0.001095 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001096 | 0.001569 | 0.001651 | 0.001347 | 0.001372 | 0.001029 | 0.002044 | 0.001677 | 0.00148 | 0.001591 | 0.002013 | 0.001674 | 0.001251 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001563 | 0.001452 | 0.002747 | 0.00318 | 0.001441 | 3.5e-05 | -0.001668 | 0.000428 | 0.001627 | -0.002234 | 0.001026 | 0.001652 | 0.001431 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00141 | 0.001151 | -0.002047 | 0.001418 | 0.001587 | 0.001552 | 0.003982 | 0.001328 | 0.001661 | 0.001358 | 0.001858 | 0.002105 | 0.001513 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.0023 | -9.9e-05 | 0.001548 | 0.00037 | -0.001954 | 0.002483 | 0.006595 | 0.00135 | 0.001202 | 0.001547 | 0.001979 | 0.001965 | 0.000809 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.45727 | 0.99699 | 0.17374 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.47724 | 0.46748 | 0.31804 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.9186 | 1.9612 | 1.9602 | 1.9342 | 1.9416 | 1.9599 | 1.9214 | 1.9526 | 1.9436 | 1.9485 | 1.9432 | 1.9638 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.8128 | 1.793 | 1.7863 | 1.7857 | 1.7909 | 1.8115 | 1.8109 | 1.7985 | 1.8114 | 1.7897 | 1.8106 | 1.7899 | 1.7885 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.7714 | 1.7689 | 1.7685 | 1.7683 | 1.7672 | 1.768 | 1.7668 | 1.7667 | 1.7672 | 1.7662 | 1.7699 | 1.7662 | 1.7669 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002167 | 0.001791 | 0.000945 | 0.001284 | 0.00068 | 0.000956 | 0.001223 | 0.001152 | 0.001062 | 0.001376 | 0.001334 | 0.001341 | 0.001381 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002094 | 0.001828 | 0.001012 | 0.001278 | Simulation 0 to 50 C fails. | 0.001166 | 0.001162 | 0.001107 | 0.001049 | 0.001347 | 0.001331 | 0.001149 | 0.001288 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.002283 | 0.001092 | 0.001274 | Simulation 0 to 50 C fails. | 0.001262 | 0.001291 | 0.000135 | 0.001176 | 0.001488 | 0.001451 | 0.001441 | 0.001516 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002075 | 0.00177 | Simulation 0 to 50 C fails. | 0.001106 | 0.000671 | 0.000538 | 0.001228 | Simulation 0 to 50 C fails. | 0.00148 | 0.001305 | 0.001292 | 0.001278 | 0.001251 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00219 | 0.002097 | 0.001082 | 0.001363 | 0.000841 | 0.001115 | 0.001142 | 0.001158 | 0.001162 | 0.001437 | 0.001433 | 0.001416 | 0.001373 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.001922 | Simulation 0 to 50 C fails. | 0.000623 | Simulation 0 to 50 C fails. | 0.001158 | 0.000672 | Simulation 0 to 50 C fails. | 0.001065 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002008 | 0.001659 | 0.000856 | 0.001212 | 0.000601 | 0.000977 | 0.001001 | 0.000956 | 0.000892 | 0.001239 | 0.001324 | 0.00118 | 0.001314 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002001 | 0.001697 | 0.000883 | 0.00114 | 0.000739 | 0.001047 | 0.001056 | 0.000993 | 0.000941 | 0.001317 | 0.001343 | 0.001347 | 0.00118 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002086 | 0.001808 | 0.001034 | 0.001218 | 0.00074 | 0.001151 | 0.00117 | 0.001069 | 0.001051 | 0.001341 | 0.001304 | 0.001304 | 0.001104 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002131 | Simulation 0 to 50 C fails. | 0.001041 | Simulation 0 to 50 C fails. | 0.000676 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.001141 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.000675 | Simulation 0 to 50 C fails. | 0.001197 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.002377 | 0.00202 | 0.001313 | 0.001258 | 0.000845 | 0.001285 | 0.001465 | 0.001362 | 0.001254 | 0.001639 | 0.00145 | 0.001554 | 0.001538 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001124 | Simulation 0 to 50 C fails. | 0.000531 | 0.000644 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001936 | 0.00175 | 0.000686 | 0.001123 | 0.000656 | 0.000984 | 0.001076 | 0.00098 | 0.000893 | 0.00127 | 0.001187 | 0.000883 | 0.00118 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.001317 | 0.001248 | 0.00076 | 0.001223 | 0.001133 | 0.001072 | 0.001032 | 0.001221 | Simulation 0 to 50 C fails. | 0.001309 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001764 | 0.001574 | 0.000771 | 0.001058 | 0.000452 | 0.000985 | 0.000932 | 0.000905 | 0.00082 | 0.000821 | 0.000796 | 0.001173 | 0.000922 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.001752 | 0.000953 | 0.001127 | 0.000696 | 0.001045 | 0.001107 | 0.001058 | 0.000975 | 0.001305 | Simulation 0 to 50 C fails. | 0.00126 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.1764 | Simulation 0 to 50 C fails. | 0.13283 | 1.8265 | 0.21875 | 0.23685 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 1.8636 | 1.547 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.9564 | 1.9563 | 1.9562 | 1.9506 | 1.9506 | 1.9504 | 1.9556 | 1.95 | 1.9498 | 1.9549 | 1.9497 | 1.9548 | 1.9492 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.7988 | 1.7987 | 1.7986 | 1.7936 | 1.7935 | 1.7981 | 1.7979 | 1.7977 | 1.7975 | 1.7974 | 1.7974 | 1.7973 | 1.797 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.7763 | 1.7762 | 1.7761 | 1.7758 | 1.7756 | 1.7755 | 1.7753 | 1.7751 | 1.7749 | 1.7749 | 1.7748 | 1.7746 | 1.7746 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001413 | 0.003456 | 0.002493 | 0.001131 | 0.000295 | 0.000884 | 0.000789 | 0.001251 | 0.00183 | 0.001118 | -4.7e-05 | -0.000416 | 0.000423 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00144 | 0.003474 | 0.002867 | 0.001172 | 0.000352 | 0.000964 | 0.000904 | 0.001302 | 0.00187 | 0.001051 | 0.000122 | -0.000406 | 0.000468 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001491 | 0.003468 | 0.003043 | 0.001253 | 0.00037 | 0.001038 | 0.000967 | 0.001302 | 0.001937 | 0.001204 | Simulation 0 to 50 C fails. | -0.000304 | 0.000463 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001356 | 0.003428 | 0.002837 | 0.001085 | Simulation 0 to 50 C fails. | 0.00083 | 0.000755 | 0.001208 | 0.0018 | 0.001084 | -1e-05 | -0.000418 | 0.000114 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001488 | 0.003546 | 0.002838 | 0.001233 | 0.000392 | 0.001343 | 0.000926 | 0.001339 | 0.001924 | 0.001277 | 0.000201 | -0.000343 | 0.000563 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.00117 | Simulation 0 to 50 C fails. | 0.000955 | Simulation 0 to 50 C fails. | 0.001123 | Simulation 0 to 50 C fails. | 0.001164 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.000481 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001268 | 0.003306 | 0.002702 | 0.000977 | 0.000197 | 0.000741 | 0.000679 | 0.001065 | 0.001667 | 0.000998 | -0.000137 | -0.000603 | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001303 | 0.003111 | 0.002792 | 0.001055 | Simulation 0 to 50 C fails. | 0.000736 | 0.000719 | 0.001138 | 0.001733 | 0.001042 | -3.4e-05 | -0.000425 | 0.000877 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001382 | 0.003452 | 0.002866 | 0.001118 | 0.000276 | 0.00087 | 0.000793 | 0.001209 | 0.001811 | 0.001118 | -8.1e-05 | Simulation 0 to 50 C fails. | 0.00067 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.000154 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.000565 | 0.000279 | Simulation 0 to 50 C fails. | 0.000574 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.001128 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00169 | 0.003412 | 0.002847 | 0.001231 | 0.000478 | 0.001165 | 0.001001 | 0.001511 | 0.001859 | 0.001365 | -0.000113 | -0.000485 | 0.000693 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 0.00374 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | 0.000442 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.00138 | 0.003374 | 0.002615 | 0.001013 | 0.000204 | 0.000761 | 0.000652 | 0.001242 | 0.001682 | 0.001029 | -8.7e-05 | -0.000491 | 0.000427 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001389 | 0.00346 | 0.002838 | 0.001244 | 0.00032 | 0.000869 | 0.000804 | 0.001219 | 0.001841 | 0.001098 | Simulation 0 to 50 C fails. | -0.000416 | 0.000147 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001144 | 0.002958 | 0.002631 | 0.00091 | 8.1e-05 | 0.000681 | 0.000605 | 0.001028 | 0.001701 | 0.001306 | -0.000106 | -0.000616 | 0.000293 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 0.001097 | 0.003413 | 0.002811 | 0.001092 | 0.00026 | 0.000798 | 0.000747 | 0.001152 | 0.001667 | 0.001035 | 6e-05 | -0.00051 | 0.00033 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Simulation 0 to 50 C fails. | 1.8318 | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Simulation 0 to 50 C fails. | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.9505 | 1.9506 | 1.956 | 1.9502 | 1.9557 | 1.9554 | 1.9556 | 1.9495 | 1.9494 | 1.9549 | 1.9547 | 1.949 | 1.9457 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.7986 | 1.7987 | 1.7985 | 1.7983 | 1.7981 | 1.7979 | 1.7976 | 1.7976 | 1.7974 | 1.7973 | 1.7972 | 1.7969 | 1.8055 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | 1.7762 | 1.776 | 1.7759 | 1.7757 | 1.7756 | 1.7754 | 1.7752 | 1.775 | 1.7747 | 1.7747 | 1.7746 | 1.7744 | 1.7743 | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error | Rebuild Error
Calcite XY Shear |  | N/mm^2 | 0.023587 | 0.00014754 | 0.027494 |  | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027495 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027495 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | 0.027494 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.0017176 | 0.0084678 | 0.00027468 | #VALUE! | #VALUE! | 0.0017177 | 0.0017177 | 0.00076327 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027494 | 0.027495 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | 0.027495 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | #VALUE! | 0.023587 | 0.023587 | 0.00028436 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | #VALUE! | 0.023587 | 0.023587 | 0.0058934 | 0.023587 | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | #VALUE! | 0.0058934 | #VALUE! | 0.023587 | 0.0058933 | #VALUE! | 0.023587 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | #VALUE! | 0.023587 | #VALUE! | 0.023587 | #VALUE! | #VALUE! | 0.023587 | #VALUE! | #VALUE! | 0.0058932 | #VALUE! | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.0058935 | #VALUE! | 0.0058935 | 0.0058934 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023586 | 0.023586 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | 0.023587 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.010819 | #VALUE! | 0.00014754 | 0.023588 | 0.00037479 | 0.00046542 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023588 | 0.017938 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023588 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | 0.023587 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.01997 | 0.019969 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.019969 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.01997 | 0.019969 | #VALUE! | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.019969 | 0.019969 | #VALUE! | 0.01997 | 0.019969 | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.019969 | #VALUE! | 0.019969 | #VALUE! | 0.01997 | #VALUE! | 0.019969 | #VALUE! | #VALUE! | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.019969 | 0.019969 | #VALUE! | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.01997 | #VALUE! | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.00024149 | #VALUE! | #VALUE! | 0.0049917 | 0.01997 | #VALUE! | 0.010066 | #VALUE! | #VALUE! | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | #VALUE! | #VALUE! | #VALUE! | 0.0049917 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.01997 | 0.019969 | 0.019969 | #VALUE! | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | 0.01997 | 0.01997 | 0.019969 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | 0.019969 | 0.01997 | 0.019969 | 0.019969 | 0.019969 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.01997 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.019971 | 0.019971 | 0.019971 | 0.019971 | 0.019971 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.019971 | 0.019971 | 0.019971 | 0.019971 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.019971 | 0.019971 | 0.019971 | 0.019971 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | 0.01997 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
SUM of stresses |  |  |  |  | 0.10474 |  | 0.10473 | 0.10473 | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 |  | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 |  | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [6 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 |  | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10474 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 |  | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10474 | [5 empty] | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [5 empty] | 0.10474 | 0.10474 | 0.10473 | 0.10473 | 0.10474 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | 0.10473 | [9 empty] | 0.0065446 | 0.032255 | 0.0010457 |  |  | 0.0065407 | 0.0065407 | 0.0029063 | [6 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10476 | 0.10474 | [6 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | [5 empty] | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | 0.10474 | [5 empty] | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090016 | [5 empty] | 0.090012 | 0.090018 | 0.090012 | 0.090012 |  | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090011 | 0.09001 | [6 empty] | 0.090018 | 0.090018 | 0.090018 |  | 0.090018 | 0.090012 | 0.0010884 | 0.090017 | 0.090012 | 0.090012 | 0.090012 | 0.090011 | [5 empty] | 0.090012 | 0.090012 |  | 0.090018 | 0.090012 | 0.022496 | 0.090017 |  | 0.090017 | 0.090012 | 0.090012 | 0.090012 | 0.090011 | [23 empty] | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090017 | 0.09001 | [6 empty] | 0.090012 |  | 0.022496 |  | 0.090012 | 0.022496 |  | 0.090012 | [9 empty] | 0.090012 | 0.090012 | 0.090017 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090017 | 0.090011 | [5 empty] | 0.090012 | 0.090012 | 0.090018 | 0.090012 | 0.090017 | 0.090012 | 0.090012 | 0.090016 | 0.090017 | 0.090012 | 0.090012 | 0.090011 | 0.09001 | [5 empty] | 0.090012 | 0.090018 | 0.090018 | 0.090012 | 0.090012 | 0.090012 | 0.090018 | 0.090018 | 0.090012 | 0.090012 | 0.090017 | 0.090011 | 0.090011 | [5 empty] | 0.090018 |  | 0.090012 |  | 0.090018 |  |  | 0.090012 |  |  | 0.022495 |  | 0.090011 | [5 empty] | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.09001 | 0.09001 | [23 empty] | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090017 | [7 empty] | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 |  | 0.090011 | [6 empty] | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090011 | 0.090011 | [6 empty] | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090012 | 0.090017 | 0.090012 | 0.090017 |  | 0.090012 | [6 empty] | 0.041293 |  | 0.00056454 | 0.090017 | 0.0014328 | 0.0017794 |  |  |  |  | 0.09002 | 0.06846 | [6 empty] | 0.090018 | 0.090018 | 0.090018 | 0.090031 | 0.090041 | 0.090042 | 0.09002 | 0.090028 | 0.090036 | 0.090017 | 0.090041 | 0.090017 | 0.090018 | [5 empty] | 0.090018 | 0.090018 | 0.090017 | 0.090018 | 0.090029 | 0.090018 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | [5 empty] | 0.090018 | 0.090018 | 0.090017 | 0.090018 | 0.090019 | 0.090019 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | 0.090017 | [5 empty] | 0.079047 | 0.079046 | 0.079046 | 0.079047 | 0.079047 | 0.079047 | 0.079048 | 0.079047 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079048 | [5 empty] | 0.079047 | 0.079047 | 0.079054 | 0.079048 | 0.079047 | 0.079047 | 0.079055 | 0.079047 | 0.079055 | 0.079048 | 0.079048 | 0.079048 | 0.079049 | [5 empty] | 0.079047 | 0.079047 | 0.079046 | 0.079054 | 0.079047 | 0.079047 | 0.07905 | 0.079049 | 0.079055 | 0.079047 |  | 0.079048 | 0.079048 | [5 empty] | 0.079053 | 0.079053 | 0.079046 | 0.079047 |  | 0.079054 | 0.079048 | 0.079055 | 0.079055 | 0.079047 | 0.079048 | 0.079048 | 0.079048 | [23 empty] | 0.079047 | 0.079047 | 0.079046 | 0.079047 | 0.079054 | 0.079054 | 0.079048 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079048 | 0.079048 | [8 empty] | 0.079047 |  | 0.079047 |  | 0.079055 |  | 0.079053 |  |  | 0.079048 | [5 empty] | 0.079047 | 0.079047 | 0.079053 | 0.079048 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079047 | 0.079048 | 0.079048 | 0.079054 | [6 empty] | 0.079047 | 0.079047 | 0.079046 | 0.079047 |  | 0.079047 | 0.079048 | 0.079047 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079048 | [5 empty] | 0.079054 | 0.079053 | 0.079047 | 0.079054 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079047 | 0.079048 | 0.079055 |  | 0.079048 | [5 empty] | 0.00095949 |  |  | 0.019765 | 0.079054 |  | 0.039851 |  |  | 0.079047 | [8 empty] | 0.079047 | 0.079047 | 0.079046 | 0.079047 | 0.079047 | 0.079047 | 0.079048 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079048 | 0.079048 | [6 empty] | 0.079047 |  |  |  | 0.019766 | [12 empty] | 0.079047 | 0.079053 | 0.079046 | 0.079047 | 0.079047 | 0.079047 | 0.079048 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079048 | 0.079048 | [5 empty] | 0.079054 | 0.079047 | 0.079055 | 0.079054 | 0.079047 | 0.079047 | 0.079048 | 0.079055 | 0.079047 | 0.079048 |  | 0.079048 | 0.079048 | [5 empty] | 0.079047 | 0.079047 | 0.079047 | 0.079047 | 0.079047 | 0.079047 | 0.079048 | 0.079047 | 0.079047 | 0.079048 | 0.079048 | 0.079048 | 0.079048 | [5 empty] | 0.079047 | 0.079047 | 0.079047 | 0.079047 | 0.079054 | 0.079047 | 0.079048 | 0.079048 | 0.079047 | 0.079055 | 0.079048 | 0.079048 | 0.079049 | [6 empty] | 0.07929 | [12 empty] | 0.079054 | 0.079055 | 0.079055 | 0.079055 | 0.079055 | 0.079055 | 0.079056 | 0.079055 | 0.079056 | 0.079055 | 0.079055 | 0.079057 | 0.079056 | [5 empty] | 0.079054 | 0.079054 | 0.079055 | 0.079055 | 0.079054 | 0.079054 | 0.079055 | 0.079055 | 0.079055 | 0.079056 | 0.079055 | 0.079055 | 0.079056 | [5 empty] | 0.079054 | 0.079054 | 0.079055 | 0.079055 | 0.079054 | 0.079055 | 0.079055 | 0.079055 | 0.079055 | 0.079055 | 0.079055 | 0.079055 | 0.079056
0.00056454 |  |  | 0.00056454
0.10476
Design Study 1
Scenarios/Iterations:
Parameter Constraint or Goal | Scenario 687
Calculated
Inner Ring Inner Dia | 36.5 | mm
Inner Ring Outer Dia | 62 | mm
Adhesive Patch D | 4 | mm
Calcite X Normal | 0.000199 | N/mm^2
Calcite Y Normal | 0.00013 | N/mm^2
Calcite XY Shear | 0.021399 | N/mm^2
Calcite YZ Shear | 8.8e-05 | N/mm^2
Calcite Z Disp | 0.13283 | mm
0.00014754
0.00056454